1. ------IND- 2010 0143 CZ- EN- ------ 20100319 --- --- PROJET

The Czech Metrology Institute (CMI), as the competent local body for stipulating metrological and technical requirements for a specified measuring device and for the stipulation of test methods for type approval and for verification of a specified measuring device pursuant to Section 14(1) of Act No. 505/1990 Coll., on Metrology, as amended, and pursuant to provisions of Section 172 et seq. of Act No. 500/2004 Coll., the Administrative Code, as amended (hereinafter the "AC"), has on 21 December 2009 commenced ex officio proceedings pursuant to Section 46 of the AC, and based on supporting documents issues this:

I.

M E A S U R E  O F  A  G E N E R A L  N A T U R E

number: 0111-OOP-C006-09

reference no. 0313/006/09/Pos.,

publication date: 21 December 2009

stipulating metrological and technical requirements for specified measuring devices, including test methods for verifying specified measuring devices:

"electronic medical and veterinary thermometers"

1   Basic definitions

For the purposes of this measure of a general nature, terminology and definitions contained in the VIM1) and the following apply:

1.1    Compact electronic thermometers

1.1.1    compact electronic thermometer: a dermal thermometer composed of a temperature probe and a display unit that are permanently joined

1.1.2    electronic medical thermometer: a thermometer used to measure body temperature in human medicine

1.1.3   electronic veterinary thermometer: a thermometer used to measure body temperature in veterinary medicine

1.1.4    thermometer with extrapolation: a device that calculates, without waiting for temperature equilibrium to be achieved, using temperature data transmission and a mathematical algorithm, the maximum temperature of a probe in contact with a body cavity, and preserves the calculated maximum temperature value, either for a specified time, or until cancelled by the user

1.1.5   thermometer without extrapolation: a device with a part or function that indicates a temperature for a required amount of time, measured by a temperature probe in contact with a body cavity, which it then displays and maintains at the maximum temperature value, either for a specified time, or until cancelled by the user

1.1.6   electronic thermometer for continuous measurement: a device composed of a display unit and connected temperature probe that continuously measures and displays body temperature

1.1.7   display unit: the part of an electronic thermometer that processes the signal output by the temperature probe and displays the temperature value

1.1.8    maximum device (of a thermometer): a part of function of an electronic thermometer that stores and displays the maximum temperature as a numerical value

1.1.9   temperature probe: the part of a thermometer used to determine body temperature and that contains a temperature sensor and other parts, including a cover, seals, internal conductors and if necessary, connectors

1.2    Electronic infrared non-contact thermometers

1.2.1   electronic infrared non-contact thermometer: an optoelectronic device capable of performing non-contact temperature measurement using heat radiated by the given location

1.2.2    electronic infrared non-contact ear thermometer: a thermometer used to ascertain the internal temperature of the human body by measuring inside the ear canal

1.2.3   electronic infrared non-contact dermal thermometer: a thermometer used to ascertain the surface temperature of the human body, most often measured on the surface of the forehead (forehead thermometer, temporal scanning thermometer, etc.)

1.2.4   ear or dermal mode: a mode in which an electronic infrared thermometer displays the temperature measured in the ear canal or on the skin of the subject; this mode allows for corrections to compensate for discrepancies, e.g. environmental conditions or emission capabilities

1.2.5   calibration mode: a mode in which an electronic infrared medical thermometer displays the measured temperature of a reference radiation source

1.2.6   estimation mode: a mode in which an electronic infrared ear thermometer displays an estimated temperature of a certain body part other than the ear canal

2   Metrological requirements

Metrological requirements are founded on requirements of Government Order No. 336/2004 Coll., laying down requirements for medical devices2), and also incorporate requirements of harmonized standards. During verification of measuring device models approved by the Metrology Act, metrological requirements extant during their introduction apply.

2.1    Compact electronic thermometers

2.1.1   Service conditions

The minimum ambient temperature range is from +10 °C to +35 °C.

2.1.2   Measurement range

The measurement range must be at least from 35.5 °C to 42.0 °C. If the measured temperature value is outside this minimum measurement rage, the thermometer must provide a visual or audio warning.

Digital increments of the displayed unit must be at most 0.1 °C.

2.1.3   Maximum allowed error

The maximum allowed error of medical thermometers in the measurement range from 35.5 °C to 42.0 °C is ± 0.1 °C in an ambient temperature range of 18 °C to 28 °C.

The maximum allowed error of veterinary thermometers under the same conditions is ± 0.2 °C.

Outside the measurement range or outside the ambient temperature range, the maximum allowed error must not be greater than double this specified value.

2.2    Electronic thermometers for continuous measurement

2.2.1   Service conditions

The minimum ambient temperature range must be from +10 °C to +40 °C, the minimum humidity range from 30% to 75% relative humidity.

2.2.2   Measurement range

The continuous measurement range must be at least from 25.0 °C to 45 °C. If the measured temperature value is outside the thermometer’s minimum measurement rage, it must provide a visual or audio warning.

Digital increments of the displayed unit must be at most 0.1 °C.

2.2.3   Maximum allowed error

The maximum allowed error of the complete thermometer in the measurement range from 25.0°C to 45.0 °C is ± 0.2 °C. The maximum allowed error of the display unit or temperature probe must not be greater than ± 0.1 °C.

For thermometers whose measuring range is greater than 25.0 °C to 45 °C, the maximum allowed error for temperatures below 25.0 °C and above 45 °C must not be greater than double the specified values.

2.3    Electronic infrared thermometers

2.3.1   Service conditions

The minimum ambient temperature range must be from +16 °C to +35 °C, the minimum humidity range up to 85% relative humidity without condensation.

2.3.2   Measurement range

2.3.2.1   An electronic infrared ear thermometer must have a temperature measuring range in all modes from at least 35.5 °C to 42.0 °C.

2.3.2.2    An electronic infrared dermal thermometer must have a temperature measuring range in all modes from at least 22.0 °C to 40.0 °C.

2.3.2.3   Digital increments of the displayed unit must be at most 0.1 °C.

2.3.3   Maximum allowed error

2.3.3.1    The maximum allowed error of ear thermometers within the extent of service conditions specified by Article 2.3.1 and in the section of the measurement range from 36.0 °C to 39.0 °C is ± 0.2 °C.

Outside the above section of the measurement range or outside the extent of service conditions, the maximum allowed error is ± 0.3 °C.

2.3.3.2   The maximum allowed error of dermal thermometers over the entire range of service conditions and over the entire measurement range is ± 0.3 °C.

3   Technical requirements

Technical requirements are founded on requirements of Government Order No. 336/2004 Coll., laying down requirements for medical devices2) and also incorporate requirements of harmonized standards. During verification of measuring device models approved by the Metrology Act, metrological requirements applicable during their introduction apply.

3.1    Compact electronic thermometers

3.1.1   Display unit

The numerals on the display must be at least four millimetres high or optically magnified to appear this high. After power is switched on, all segments and symbols on the display unit must be activated for at least one second.

The display unit must include a means of automatic checking whose accuracy is equal to or better than the display unit’s maximum allowed error.

3.1.2   Power source

Compact electronic thermometers must automatically provide a visual or audio warning that the voltage of the power source is outside the limits stipulated by the manufacturer.

3.2   Electronic thermometers for continuous measurement

3.2.1   Display unit

The numerals on the display must be at least four millimetres high or optically magnified to appear this high. The numerical values on the display unit must refresh at least every 10 seconds. After power is switched on, all segments and symbols on the display unit must be activated for at least two seconds.

The display unit must include a means of automatic checking with an accuracy equal to or better than the display unit’s maximum allowed error stipulated by the manufacturer.

3.2.2   Temperature probe

If the temperature probe, sans cover, is immersed into an electrically conductive liquid, its insulation must be such that the change in displayed temperature is at most ± 0.02 °C.

The temperature probe must meet the requirements stipulated in article 2.2.3 even after cleaning, disinfection or sterilisation in accordance with the manufacturer’s specifications.

3.2.3   Power source

When powered from the mains power supply, the temperature displayed by the thermometer must not change when the power supply voltage fluctuates by ± 10 % and the frequency by ± 2 % compared to nominal values.

When being powered by a battery or auxiliary power source, the thermometer must provide a visual or audio warning that the power supply voltage is equal to or less than the value stipulated by the manufacturer.

3.3   Electronic infrared thermometers

3.3.1    Infrared thermometer types

An electronic infrared ear or dermal thermometer must function in ear or dermal mode whose measurement conditions must be stipulated by the manufacturer.

A calibration mode must be available either by switching the measuring device directly into this mode, or by switching from ear or dermal mode, or it can be identical to them. If estimation modes are available (e.g. marrow, rectal, oral), the values displayed on the display unit must be labelled in an unmistakeable way as estimated values.

3.3.2   Display unit

The numerals on the display must be at least four millimetres high or optically magnified to appear this high. After power is switched on, all segments and symbols on the display unit must be activated for at least two seconds.

An electronic infrared thermometer must have automatic monitoring of self-checking functions and proper operation must be indicated by a suitable display on the display unit. The thermometer must display its current mode.

An electronic infrared thermometer must provide a visual warning or not display temperature data if the measured temperature, ambient temperature or power supply voltage are outside the range stipulated by the manufacturer.

3.3.3   Temperature probe

The temperature probe must meet the requirements stipulated in Article 2.3.3.1 or 2.3.3.2 even after cleaning, disinfection and sterilisation according to the manufacturer's specifications.

If the manufacturer requires that a temperature probe cover be in place while using the electronic infrared thermometer, it must not display a temperature indication without the cover in place, or the display unit must show information that the probe cover must be replaced prior to measuring body temperature. When placing the cover on the temperature probe, its physical integrity must be preserved.

If the thermometer is designed for use with protective temperature probe covers, the thermometer along with the probe cover must meet the requirements stipulated by this measure.

3.3.4   Power supply

When powered from the mains power supply, the temperature displayed must not change when the power supply voltage fluctuates by ± 10 % and the frequency by ± 2 % compared to nominal values.

When being powered by a battery or auxiliary power source, the thermometer must provide a visual or audio warning, or it must not display temperature, if the voltage falls outside the range stipulated by the manufacturer.

4   Measuring device labelling

4.1   In general

The thermometer must have information provided by the manufacturer and graphical markings for designation of medical aids. The thermometer must also be labelled with the CE mark and the number of the notified subject that assessed its conformance. In the case of thermometers introduced in accordance with the Metrology Act, the thermometer must be labelled with a type approval mark.

All labels and markings must be easily visible and readable under normal working conditions, permanent, and must not hinder the reading of the measuring device's data.

4.2    Labels and markings on compact electronic thermometers

Aside from labelling required according to Article 4.1, the measuring device must be labelled with at least the following information:

the manufacturer’s name,

model or type name,

serial number,

measuring range and a "°C" measuring unit marking associated with the numerical value (if not indicated on the display),

orientation or position during use, if this affects the measurement result,

information that the temperature value displayed is the result of a calculation,

information on the specific temperature measurement range, if the overall measurement range is greater,

body location (e.g. mouth or anus), if there is any limitation, e.g. for thermometers with extrapolation,

the warning "Not water-resistant" or equivalent declaration if the thermometer does not meet requirements for contact with water.

4.3    Labels and markings on electronic thermometers for continuous measurement

Aside from labelling required according to Article 4.1, the measuring device must be labelled with at least the following information:

- the manufacturer’s name,

- model or type name,

- serial number,

- measuring range and a "°C" measuring unit marking associated with the numerical value (if not indicated on the display),

- orientation or position during use, if this affects the measurement result,

- for the temperature probe, the location on the body where temperature should be measured, if there is a restriction (for example the mouth or anus),

- indication of the specific measurement range, if it is greater than the minimum measurement range (see Article 2.2.3).

4.4    Labels and markings on electronic infrared thermometers

Aside from labelling required according to Article 4.1, the measuring device must be labelled with at least the following information:

- the manufacturer’s name,

- model or type name,

- serial number,

- year and month of manufacture and initial calibration, if different

- measuring range and a "°C" measuring unit marking associated with the numerical value (if not indicated on the display),

- measurement location on the body,

- estimation modes, if not indicated on the display,

- whether protective probe covers are required.

4.5   Special information provided by the manufacturer

For electronic infrared thermometers, the manufacturer must provide information on clinical accuracy, especially information on temperature deviation of the body temperature measurement location, clinical systematic error, its standard deviation and the derivation of these estimates.

He must also provide information on the sequence and effect of the thermometer’s self-checking functions.

5    Measuring device type approval

These measuring devices are not subject to type approval in accordance with Act No. 505/1990 Coll., on metrology, as amended.

6   Initial verification

The initial verification procedure3) is identical to the subsequent verification procedure in accordance with Chapter 7.

7   Subsequent verification

The following tests are performed during subsequent verification of electronic thermometers:

- a visual inspection,

- an accuracy test.

7.1   Visual inspection

The following is checked during a visual inspection:

whether the thermometer submitted for verification conforms to the approved type or with a design for which conformance was declared during market introduction,

whether it is not physically damaged or does not exhibit defects visible to the naked eye,

whether the lens of an infrared thermometer’s outer optical assembly does not exhibit defects,

whether it is properly labelled.

If the thermometer does not pass external inspection, it is not tested further.

7.2    Compact electronic thermometer accuracy test

7.2.1   Test equipment

The following test equipment is used:

a reference thermometer with a display device resolution of 0.01 °C and an extended measurement uncertainty less than or equal to ± 0.02 °C (extension coefficient k = 2) to determine the temperature of the liquid bath,

a liquid bath with a volume of at least 5 l for a minimum liquid temperature of +30 °C to +45 °C, with temperature regulation and a mixing device, that ensure temperature stability of less than or equal to ± 0.02 °C and a temperature field homogeneity less than or equal to ± 0.01 °C at an arbitrary temperature.

7.2.2   Test method

The accuracy test involves comparing values measured by the thermometer with temperature values on the reference thermometer in a liquid bath at an ambient reference temperature of (23 ± 5) °C and relative humidity of (50 ± 20) % with the device operating within the specified power supply voltage range.

Thermometers having a measurement interval range of ≤ 10 °C are tested at minimally two points, and those with a measurement range of > 10 °C at minimally three points that include 37 °C and 41 °C for medical thermometers and 38 °C and 42 °C for veterinary thermometers.

The difference between the measured temperature value and the reference temperature value must meet requirements specified in Article 2.1.3 at every measurement point.

7.3    Continuous-measurement electronic thermometer accuracy test

7.3.1   Test equipment

The following test equipment is used:

- a reference thermometer with a display device resolution of 0.01 °C and an extended measurement uncertainty less than or equal to ± 0.02 °C (extension coefficient k = 2) to determine the temperature of the liquid bath,

- a liquid bath with a volume of at least 5 l for a minimum liquid temperature of +30 °C to +45 °C, with temperature regulation and a mixing device that ensure temperature stability of less than or equal to ± 0.02 °C and a temperature field homogeneity less than or equal to ± 0.01 °C at an arbitrary temperature,

- a temperature probe simulator with extended measurement uncertainty no greater than 0.01 °C (calculated for an extension coefficient of k = 2) according to information from the manufacturer,

- a temperature probe tester to transform the measured physical characteristics of the probe to temperature values by measuring changes in these characteristics related to temperature with an extended measurement uncertainty of this physical characteristic that in the measurement range is no greater than 0.01 °C (calculated for an extension coefficient of k = 2) according to information from the manufacturer.

7.3.2   Test method

An accuracy test for continuous-measurement electronic thermometers may be performed either as a test of the complete thermometer, or as independent tests of the display unit and the temperature probe.

7.3.2.1    Test of the complete thermometer

The accuracy test involves comparing values measured by the thermometer with temperature values on the reference thermometer in a liquid bath at an ambient reference temperature of (23 ± 5) °C and relative humidity of (50 ± 20) % with the device operating within the specified power supply voltage range.

Thermometers having a measurement interval range of ≤ 10 °C are tested at minimally three points, and those with a measurement range of > 10 °C at minimally five points that include 36 °C, 37 °C and 41 °C for medical thermometers.

The difference between the measured temperature value and the reference temperature value must meet requirements specified in Article 2.2.3 at every measurement point.

7.3.2.2   Test of part of the thermometer

Display unit characteristics are tested using a temperature probe simulator according to 7.3.1 c).

The difference between the temperature displayed on the display unit and the corresponding simulated temperature values must meet display unit requirements according to Article 2.2.3.

When testing only the probe, it is immersed in a calibration temperature bath like for the test of the complete thermometer. The temperature probe is connected to a tester according to 7.3.1 d). Each temperature probe value obtained in this manner is compared with the value displayed by the reference thermometer in the bath.

The difference between the measured temperature value and the reference temperature value must meet requirements specified in Article 2.2.3 at every measurement point.

7.4    Electronic infrared thermometer accuracy test

7.4.1   Test equipment

The following test equipment is used for the accuracy test:

- a reference thermometer with a display device resolution of 0.01 °C and an extended measurement uncertainty less than or equal to ± 0.03 °C (extension coefficient k = 2) to determine the temperature of a black body,

- a liquid bath with a volume of at least 3 l for a minimum liquid temperature of +30 °C to +45 °C, with temperature regulation and a mixing device that ensure temperature stability of less than or equal to ± 0.02 °C and a temperature field homogeneity less than or equal to ± 0.01 °C at an arbitrary temperature,

- a black body in the shape of a cavity, modified for use in a liquid bath, with the ability to set black body temperature using a reference thermometer with an evaluation device; the temperature of the black body must be specified with an extended measurement uncertainty not exceeding ± 0.07 °C (extension coefficient k = 2),

- an aperture to decrease the size of the opening into the cavity of the black body that permits proper positioning of the tested thermometer in the black body.

7.4.2   Test method

The accuracy test of the thermometer, either alone or with a specified temperature probe cover (if called for by the manufacturer) must be performed in calibration mode; if this mode is not available, then in ear or dermal mode.

Thermometers are tested at a reference temperature of (23 ± 5) °C and relative humidity of (50 ± 20)% by measuring the radiation of a black body with an aperture immersed in a liquid bath.

Thermometer accuracy is determined by comparing the values measured by the thermometer with the reference temperature value of the black body in the liquid bath. The black body temperature is measured by the thermometer being tested according to the procedure recommended by the manufacturer.

The accuracy of ear thermometers is usually tested at three points at temperatures of 36 °C, 37 °C and 41 °C set with an accuracy of ± 0.5 °C.

The accuracy of dermal thermometers is usually tested at three points at temperatures of 23 °C, 30 °C and 38 °C set with an accuracy of ± 1.0 °C.

The difference between the measured temperature value and the reference temperature value must meet requirements specified in Article 2.3.3 at every measurement point.

8   Notified standards

To specify metrological and technical requirements for measuring devices and to specify verification test methods stemming from this measure of a general nature, the CMI notifies Czech technical standards, other technical standards or technical documents of international or foreign organizations or other technical documents containing detailed technical requirements (hereinafter "notified standards"). The CMI shall publish a list of these notified standards, with their assignation to relevant measures, along with the measure of a general nature, in a manner available to the public (on the website www.cmi.cz).

Compliance with notified standards or parts thereof is considered to be, to the extent and under the conditions stipulated by a measure of a general nature, to be compliance with those requirements stipulated by this measure to which these standards or parts thereof apply.

II.

J U S T I F I C A T I O N

The CMI issues, towards the implementation of Section 24c of Act No. 505/1990 Coll., on metrology, as amended, this measure of a general nature, stipulating metrological and technical requirements for specified measuring devices and test methods for verification of these specified measuring devices.

In Decree No. 345/2002 Coll., specifying measuring devices for mandatory verification and measuring devices subject to type approval, as amended, the annex "List of Specified Measuring Device Types" classifies electronic medical and veterinary thermometers under item 3.1.1 as measuring devices subject to verification.

The CMI thus issues, towards the implementation of Section 24c of Act No. 505/1990 Coll., on metrology, as amended, for the measuring device type specified as "electronic medical and veterinary thermometers", this measure of a general nature, stipulating metrological and technical requirements for electronic medical and veterinary thermometers and test methods for verification of these specified measuring devices.

III.

E F F E C T I V E  D A T E

This measure of a general nature takes effect on the fifteenth day following the day of its publication (Section 24d of Act No. 505/1990 Coll., on metrology, as amended).

This measure of a general nature shall be published for a time span of 15 days.

# Footnotes

1)   International Vocabulary of Metrology - Basic and General Concepts and Associated Terms (VIM)

2)  This government order implements Council Directive 93/42/EEC of 14 June 1993 concerning medical devices, as amended.

3) Initial verification applies only to measuring devices that at that time have a valid type approval certificate, and to measuring devices following repair.