1. ------IND- 2010 0348 CZ- EN- ------ 20100625 --- --- PROJET

The Czech Metrology Institute (CMI), as the competent local body for stipulating metrological and technical requirements for a specified measuring device and for the stipulation of test methods for type approval and for verification of a specified measuring device pursuant to Section 14(1) of Act No. 505/1990 Coll., on Metrology, as amended, and pursuant to provisions of Section 172 et seq. of Act No. 500/2004 Coll., the Administrative Code, as amended (hereinafter the "AC"), has on 26 February 2010 commenced ex officio proceedings pursuant to Section 46 of the AC, and based on supporting documents issues this:

I.

M E A S U R E  O F  A  G E N E R A L  N A T U R E

number 0111-OOP-C008-10

reference no. 0313/001/10/Pos.,

publication date: 26 Feb. 2010

stipulating metrological and technical requirements for specified measuring devices, including test methods for verifying specified measuring devices:

"alcohol meters used to measure the volume of alcohol produced"

1   Basic definitions

For the purposes of this measure of a general nature, terminology and definitions contained in the VIM1) and the following apply:

1.1   Alcohol meter: a meter intended for the measurement of the amount of alcohol produced in commercial distilleries and grower distilleries, and to which requirements of special legislation apply; 2), 3)

1.2   Alcohol: pure ethanol (ethyl alcohol) or a mixture of ethanol and water;

1.3    Volume concentration of an ethanol-water mixture: the ratio of the volume of pure ethanol contained in a mixture at 20 °C to the total mixture volume at the same temperature;

1.4   Mass concentration of an ethanol-water mixture: the ratio of the mass of pure ethanol contained in a mixture to the total mixture mass;

1.5   Drum meter: a volume meter whose main part is a rotary drum with a horizontal axis of rotation that is composed of several measuring chambers, and which is put into motion by the effect of the weight of liquid (alcohol) as individual chambers alternately fill and empty;

1.6   Sampling apparatus: an apparatus for the automatic sampling and collection of samples of the alcohol being measured that permit the independent determination of the average volume concentration of alcohol flow;

1.7   Corrective apparatus: an auxiliary apparatus that corrects measured volume values according to alcohol flow density and temperature;

1.8   Conversion apparatus: an auxiliary apparatus that automatically determines the volume concentration of alcohol from the measured density and temperature of alcohol flow;

1.9   Backup meter: an apparatus for measuring alcohol volume if alcohol flow rate exceeds the upper limit of the measurement range of a drum meter, or in case of undue drum stoppage;

1.10    Sampling apparatus collection vessel: a vessel for the collection of alcohol samples;

1.11   Insertion vessel: a vessel located near collection vessels that serves to assess changes to alcohol concentration in collection vessels;

1.12   Surge vessel: a vessel that captures alcohol samples if alcohol outflow is blocked;

2   Metrological requirements

Metrological requirements are based on the requirements of recommendation OIML R 117-1 "Dynamic measuring systems for liquids other than water. Part 1: Metrological and technical requirements4)";

2.1   Service conditions

Service conditions are specified by the meter manufacturer; the ambient temperature range for service must be at least (0 to 50) °C;

2.2   Measurement range

The measurement range is specified by the meter manufacturer and is given by the maximum permissible alcohol flow rate through the meter;

2.3   Maximum permissible error

2.3.1    Maximum permissible flow volume error

Maximum permissible flow volume errors for alcohol drum meters is given in Table 1 by meter size

Table 1 – Maximum permissible flow volume error

| Meter type | Flow dm3/min | Maximum permissible error % |
| --- | --- | --- |
| Meters with drum capacity less than or equal to 4 dm3 (low-capacity meters) | up to 1 1 to 2 above 2 | ± 0.2 ± 0.25 ± 0.3 |
| meters with drum capacity greater than 4 dm3 (high-capacity meters) | up to 5 5 to 10 above 10 | ± 0.2 ± 0.25 ± 0.3 |
| backup meters | – | ± 0.3 |

2.3.2    Maximum permissible volume errors of structural elements

2.3.2.1    The difference between the largest and smallest chamber in terms of nominal drum chamber volume must not exceed ± 0.2 %;

2.3.2.2   The error of n volumes of sampling apparatus dippers must be smaller than ± 10 % of the n-multiple of the nominal volume of one sampling apparatus dipper, where n is the number of drum chambers;

2.3.2.3   The volume error of collection vessels must be less than ± 5 %. Volume error of insertion vessels must be smaller than -5 %. Volume error of surge vessels must be less than -2.5 %;

2.3.3    Maximum permissible errors for measuring system elements

2.3.3.1   The peak thermometer for indication that the permissible temperature range has been exceeded must have an error of less than ± 1 %;

2.3.3.2    The maximum permissible error of a connected temperature sensor/transducer is ± 0.3 °C;

2.3.3.3    The maximum permissible error of a density sensor/transducer for temperature conversion is ± 3.0 kg/m3; for volume/mass conversion and vice versa, the greatest permissible error is ± 0.6 kg/m3;

2.4    Measuring system classification (accuracy class)

For alcohol meters (drum type), an accuracy class of 0.5 applies;

3   Technical requirements

Technical requirements are based on the requirements of recommendation OIML R 117-1 "Dynamic measuring systems for liquids other than water. Part 1: Metrological and technical requirements4)";

3.1   Meter types

An alcohol meter used to measure the volume of alcohol produced is:

a drum meter with a sampling apparatus that is classified according to drum volume as:

meters with drum volume less than or equal to 4 dm3 (low-volume meters),

meters with drum volume greater than 4 dm3 (high-volume meters),

a drum meter with a conversion apparatus for determining volume concentration of alcohol.

3.1.1   Measuring drum

The volume of individual measuring chambers, i.e. the minimum amount measurable by the drum meter, must be 1 × 10n, 2 × 10n, or 5 × 10n litres, where is a positive or negative number or zero. Drum chambers must be equal in size and must have air relief valves.

The axis of the drum must be horizontal when in its operation position, and the meter must have a horizontal level indicator.

The measuring drum and indicator apparatus must be prevented from operating in reverse.

3.1.2   Sampling apparatus

The ladles of the sampling apparatus that sample alcohol from the main or backup chambered drum must be designed in such a way that they can be replaced with ladles of different volumes. For purposes of drum balancing, ladles from each set must have the same nominal weight.

The nominal volume of ladles is 1 cm3, 2.5 cm3 or 5 cm3.

The sampling apparatus must have two collection vessels with sufficient bottom slope and drainage taps.

An insertion vessel must be located near each collection vessel.

If the drained volume is subjected to special or separate treatment, the measuring apparatus must be configured in such a way that the drained volume is not present in the drum meter's indication.

3.1.3   Conversion apparatus

The conversion apparatus automatically calculates and displays the volume concentration of alcohol in the flow from its density and temperature using alcoholometric tables (see OIML R 22 – International alcoholometric tables5)) stored in its memory for a given measurement range.

The reference temperature for alcohol measurement is 20 °C.

It must be possible to calibrate the connected conversion apparatus that measures temperature and density.

3.2   Display apparatus

The display apparatus may be mechanical or electronic, and must be impossible to zero.

One division on the meter scale must equal the volume of a measuring chamber.

The maximum range of the entire display apparatus must be of the form 1 × 10k, 2 × 10k or 5 × 10k units of volume, where k is a positive whole number.

The measuring and display apparatus must be connected mechanically or with a cable.

3.3   Auxiliary apparatus

The measuring system includes the following auxiliary apparatus:

a temperature probe/transducer for determining the temperature of the measured alcohol, located in front of the alcohol's entrance into the drum;

an alcohol density sensor/transducer or spirit gauge located in front of the alcohol's entrance into the drum.

3.4    Failure prevention and meter safety

3.4.1    Unacceptable operating conditions and failures

A drum meter must, either via a special apparatus included in the meter, prevent the occurrence of or using an alarm draw attention to the following unacceptable operating conditions or failures:

flow rate exceeding the maximum permissible alcohol flow rate through the meter;

a barrier to free flow;

drum overflow due to a failure to rotate;

a temperature outside the permitted range;

unacceptable heating of a sample.

3.4.2    Emergency equipment and safety features

3.4.2.1    A high-volume meter must have a backup chambered drum with an independent display. Backup meter activity must be signalled both visibly and audibly.

3.4.2.2    A low-volume meter must have a visible indication of alcohol flow through the drum exceeding its maximum, or in case of undue drum stoppage.

3.4.2.3    There must be an apparatus to take off alcohol vapours at the location where the alcohol enters the drum.

3.4.2.4    The meter must have a peak thermometer to indicate that the highest permissible temperature has been exceeded.

3.4.2.5    The meter must have a surge vessel.

3.4.2.6    The meter must have an apparatus that prevents air from building up in the drainage pipes.

3.4.2.7    The housing containing all meter components must have a glass window for visual inspection of alcohol flow and to monitor alcohol volume concentration and temperature meters.

3.4.2.8    A high-volume meter must be protected from intentional application of heat with a protective sheath that covers the meter housing except for the glass window, nameplate and main counter display.

3.5   Material

The meter must be manufactured from material that is permanently alcohol resistant and is not harmful to health.

Individual parts of the meter must be manufactured in such a way that deformations due to temperatures within the meter's operating range that would change the measuring volume of the meter component, and thus influence its accuracy, shall not occur.

3.6    Power

Meters that use an electronic conversion apparatus must be powered in a manner suitable for an explosive environment in accordance with applicable regulations.

3.7   Resistance to external influences

Meters are only intended for measurement in buildings, i.e. in an H2 climactic environment (see also OIML R 117-1).

Meters must be rigidly fastened to their base – an M1 mechanical environment (see also OIML R 117-1).

3.7   Electromagnetic environment

For meters with connected electronic transducers and conversion apparatus, an electromagnetic class of E1 is defined (see also OIML R 117-1).

3.8   Installation requirements

To ensure meter accuracy and at the same time its on-site verification, the following installation requirements must be met:

the meter must be installed with the axis of the measuring drum in a horizontal position and must be prevented from moving and relocation;

the meter must be protected from shocks and vibrations;

the meter outflow pipe must be located at a height that permits prescribed tests to be performed;

the meter must be freely accessible from all sides, and must be located sufficiently far from sources of heat.

3.9    Meter security and protection from fraud

Meters must contain protective devices that can be sealed in such a manner that both before and after the meter's proper installation, there is no way that the meter can be taken apart or modified without breaking the seal or protective devices.

3.10   Transport and storage

During transport and long-term storage, certain removable parts must be removed, and some must be locked in place, according to manufacturer instructions.

4   Meter labelling

4.1   Meter labelling

Meters must be clearly and visible labelled with the following information:

the manufacturer's name or logo;

the meter model;

serial number and year of manufacture (the serial number must also be permanently placed on all main and some secondary parts);

meter design approval label;

nominal drum volume;

nominal drum chamber volume, i.e. the minimum measurable amount;

maximum permissible inflow;

operating temperature range;

the counter must list the measurement unit;

the range of densities or volume concentrations (for conversion apparatus).

All data, specifications, information, etc. required in the articles of this measure must be listed in the meter's operating manual.

4.2    Location of official label

The meter must have locations prepared for official labels. Locations for official and security labels are stipulated in the meter's type approval certificate.

The location for the main official label is on the meter's nameplate.

The meter's housing and protective sheath must be designed to allow the meter to be sealed off with valves.

5   Meter type approval

5.1   In general

the type approval process for an alcohol meter includes the following tests:

external inspection;

leak test;

accuracy test;

volume test for individual measuring chambers of the drum;

resolution test;

accuracy test of the backup meter;

accuracy test of structural elements;

indicator check;

test of the unit's conversion apparatus (if connected);

a test of the resistance of the conversion apparatus including sensors to disruptive influences of the surrounding environment (if connected).

5.2   External inspection

The following is assessed during an external inspection of an alcohol meter:

that prescribed technical documentation is complete;

that metrological and technical characteristics specified by the manufacturer in documentation conform to requirements of this regulation listed in sections 2 and 3;

that the alcohol meter is complete and in accordance with prescribed technical documentation;

that the software version of the electronic conversion apparatus is the same as the version specified by the manufacturer.

5.3   Functional tests

5.3.1   Test equipment

The following test equipment is used:

- graduated reference vessels and graduated cylinders of suitable volumes with an accuracy at least three times better than the greatest permissible error of the meter being tested;

- a pipette;

- a base for two graduated vessels with adjustable screw legs for horizontal levelling and a spirit level;

- a thermometer with 0.1 °C increments with a (0 to 50) °C measuring range;

- a stopwatch with 0.1 s increments;

a set of spirit gauges for a range of (0 to 100) % by volume with 0.1%, optionally also 0.2%.

A sufficient amount of alcohol to perform the tests, with a volume concentration of around 50% for low-volume meters and around 90% for high-volume meters.

5.3.2   Leak test

A leak test is performed on the drum after it is removed from the meter, cleaned and dried. The drum is suspended so that the lower chamber can be entirely filled with alcohol. As the drum is turned, the individual chambers fill with alcohol. Each chamber is tested for 5 minutes for leaks in its corners and walls.

A leak test is performed 5 minutes after being filled with alcohol for:

the backup meter;

the spirit gauge reservoir connection to the meter housing.

A leak test is performed on collection and insertion vessels and their drainage valves when new and repaired meters are being tested, 3 hours after being filled with alcohol.

If a leak is found, the meter is not subjected to further tests pursuant to this measure.

5.3.3   Accuracy test

An accuracy test for the meter measuring the volume of alcohol that has flowed through the measuring drum is performed using one of the following methods:

alternate filling and emptying of two graduated 20 dm3 reference vessels;

measurement of the amount of alcohol that has flowed through gathered in an auxiliary metal barrel using reference graduated vessels for flow rates greater than 10 dm3/min;

alternate filling and emptying of a double volume apparatus whose counter directly records the number of measurements or the total amount that has flowed through directly.

The following principles must be observed when using reference vessels:

the reference vessel must be placed in a horizontal position;

the reference vessel must be pre-wetted;

when filling the vessel, the exact volume is set by the lowest point of the liquid meniscus touching the upper edge of the graduation marking, and must be read at the level of the liquid's surface to prevent parallax error;

a drip time of 30 s must be observed after the liquid has been poured out of the vessel.

The total volume flow indicated by the meter must not be reduced by the volume taken by the sampling apparatus, and so the sampling apparatus must be disconnected during the test.

After the alcohol outflow from the meter has been disconnected from the fixed outflow pipe, alternately fill the graduated vessels, whose nominal volume is always adjusted after filling with a pipette. The number of vessels filled to their nominal volume is recorded, as are measured volume surpluses or additions.

The smallest measured amount must be equal to ten times the volume of the drum.

Measurement of the accuracy of volume that has flowed through the drum must be performed at flow rates listed in Table 2.

Table 2

| Meter | Flow rate (dm3/min) |
| --- | --- |
|  | Q1 | Q2 | Q3 |
| low-volume | 0.5 to 1 | 1 to 2 | 2 to 3 |
| high-volume | 3 to 5 | 5 to 10 | 10 to 15 |

For each 20 dm3 that have flowed through, the following must be measured:

the alcohol temperature t1 at the input to the tested meter and alcohol temperature t2 in the filled reference graduated vessel (after stirring contents); both temperatures are recorded and sued to calculate median temperatures 1 and 2 ;

the apparent volume concentration of alcohol in the filled reference graduated vessel (spirit gauge measurement).

The following calculations must then be performed:

the true volume concentration of alcohol from the apparent concentration measured by the spirit gauge;

correction of data from the alcohol meter to the temperature of alcohol in the graduated vessel.

In order to eliminate alcohol temperature changes, the volume given by the meter display and the volume measured by the graduated vessels must be recalculated.

Meter error in % is calculated as follows:

where 	V1	volume that has flowed through the meter in dm3;

V2	volume of alcohol determined by reference graduated cylinders, in dm3;

ε	measurement error of volume that has flowed through the tested meter, in %;

ρ1	alcohol density at temperature t1, in kg/m3;

ρ2	alcohol density at temperature t2, in kg/m3;

Q	flow rate, in dm3/min;

τ	test time, in min.

Measurement error ε must not be greater than values listed in Table 1 for the appropriate meter size.

5.3.4    Volume test of individual drum measuring chambers

The test is performed by pouring the chamber contents into a graduated cylinder or reference graduated vessel having a volume of:

5 l for large-volume meters;

1 l for small-volume meters.

When the volume of individual drum chambers is being tested, the first tested chamber is measured once more after all other chambers have been testing. Its volume is determined from the second measurement.

The difference in volume of the largest and smallest chamber must meet the provisions of section 2.3.2.1.

5.3.5   Resolution test

Prior to the test, the entire drum must be pre-wetted and then put into such a position that the edge of one of the outflow openings is opposite the mark on the spirit gauge reservoir.

At a minimum flow rate of Qmin the drum must start moving before the alcohol begins to flow out of the outflow channel.

The backup meter is tested in the same way.

5.3.6   Backup measuring apparatus accuracy test

The following must be checked prior to the test:

regular movement of the apparatus;

correct connection of the overflow groove;

correct engagement and movement of the lever system and alarm functionality.

A test of the accuracy of measurement the volume of alcohol that has flowed through the backup drum is performed according to section 5.3.3 at a flow rate of (5 to 10) dm3/min.

The error in measurement of the backup drum volume must not be greater than the values listed in Table 1 for the relevant meter size.

5.3.7    Structural element accuracy test

5.3.7.1    Sampling apparatus accuracy test

Determining volume errors of all sampling apparatus sets (ladles) is done during one drum rotation. The volume of alcohol sampled by n ladles, where n is the number of drum chambers, is captured for each sampling apparatus set in a graduated cylinder with a spout

The error of n ladle volumes must meet the requirements of Section 2.3.2.2.

5.3.7.2    Accuracy test of other structural elements

Tests for correct collection, insertion and surge vessel volumes are performed after pre-wetting by measuring volume using graduated cylinders and pipettes.

Volume errors of collection, insertion and surge vessels must meet the requirements of Section 2.3.2.3.

5.3.8    Display check

The check consists of:

checking that gearing used (after removing protective covers) is in accordance with meter documentation;

checking that the block preventing the display from being rolled back works properly.

5.3.9   Conversion apparatus test

If the alcohol meter has a built-in or auxiliary conversion apparatus, which uses alcoholometric tables or connected temperature and density sensors and performs conversion of the volume of a mixture of ethanol and water to normal conditions, conversion to mass, and calculates volume or mass concentrations of ethanol, this conversion apparatus is tested from the perspective of its functionality and form the perspective of storage of alcoholometric table data in the conversion apparatus.

If the conversion apparatus receives input from connected measuring devices, such as a thermometer or densitometer for correction or conversion of measured data, they must be included in the type approval of the alcohol meter.

5.4    Tests of conversion apparatus resistance to disruptive environmental influences

Of the tests listed below, those that are relevant to the design and specified conditions of use of the conversion apparatus are performed.

5.4.1    Tests of resistance to climactic influences

5.4.1.1    Limit temperatures test

Resistance to ambient temperature limits is tested with the conversion apparatus switched on:

- with dry heat at a temperature of 55 °C for 2 hours;

- with cold at a temperature of -10 °C or -25 °C (according to manufacturer requirements) for 2 hours.

After the specified time has passed, an accuracy test is performed at the given limit temperature at three points of the alcohol meter's measurement range. Measurement errors must not exceed the greatest permissible error according to Section 2.3.1.

5.4.1.2    Humidity test

Resistance to humidity is tested with the conversion apparatus switched off by cycling wet heat in two 24 hr cycles at a maximum temperature of 55 °C and a relative humidity greater than 95 %.

After both cycles have completed and restoration of operating conditions, measurement errors must not exceed the maximum permissible error according to Section 2.3.1.

5.4.2    Limit power voltages test

Resistance to limit power voltage values are tested with the conversion apparatus switched on. For AC these limits are Unom – 15 % and Unom + 10 %, where Unom is the nominal power voltage.

Measurement errors must not exceed the greatest permissible error according to Section 2.3.1 at limit power voltage values.

5.4.3    Electromagnetic compatibility (EMC) tests

5.4.3.1    Resistance to short AC voltage dips, short interruptions, and slow voltage changes

Resistance to short AC voltage gips, short interruptions and slow voltage changes are tested with the conversion apparatus switched on at all AC power inputs by applying a voltage dip:

of 100% of Unom for 0.5 of an AC voltage cycle;

of 100% of Unom for 1 AC voltage cycle;

of 20% of Unom for 25 AC voltage cycles.

The test must be performed for at least one flow rate, with simulated input signals being allowed. No serious errors may occur during this test.

5.4.3.2    Resistance to rapid transient electrical phenomena/impulse groups on power voltage

Resistance to rapid transient electrical phenomena/impulse groups is tested with the conversion apparatus switched on at a test voltage of ± 1 kV on power terminals with at least 10 positive or negative randomly phased transient phenomena/impulse groups.

The test must be performed for at least one flow rate, with simulated input signals being allowed. No serious errors may occur during this test.

5.4.3.3    Resistance to electrostatic discharge

Resistance to electrostatic discharge is tested with the conversion apparatus switched on, preferably with a contact discharge of 6 kV applied to a conductive section of the meter or with an 8 kV air discharge applied to coupling plates near the meter.

The test must be performed for at least one flow rate, with simulated input signals being allowed. No serious errors may occur during this test.

5.4.3.4    Resistance to rapid transient electrical phenomena/impulse groups on signal, data and control wiring

Resistance to rapid transient electrical phenomena/impulse groups is tested with the conversion apparatus switched on at a test voltage of ± 0.5 kV on signal terminals with a repeat frequency of 5 kHz. Interference is applied to all meter inputs and outputs.

The test must be performed for at least one flow rate, with simulated input signals being allowed. No serious errors may occur during this test.

5.4.3.5    Resistance to electrical shocks on signal, data and control wiring

Resistance to shock impulses is tested with the conversion apparatus switched on by applying at least three positive and three negative voltage shocks:

asymmetrical line for line – line 0.5 kV;

asymmetrical line for line – ground 1.0 kV;

symmetrical line for line – ground 1.0 kV.

The test must be performed for at least one flow rate, with simulated input signals being allowed. No serious errors may occur during this test.

5.4.3.6    Resistance to shock impulses on power voltage

Interference is composed of three positive and three negative voltage impulses. For AC, pulses must be applied synchronously with AC power for phase shifts of 0°, 90°, 180° and 270°. Voltage amplitude is 1.0 kV for line – line and 2.0 kV for line – ground.

The test must be performed for at least one flow rate, with simulated input signals being allowed. No serious errors may occur during this test.

5.4.3.7    Resistance to a radiated high-frequency electromagnetic field

Resistance to a radiated high-frequency electromagnetic field is tested with the conversion apparatus switched on. The conversion apparatus is subjected to an electromagnetic field in the 80 MHz to 2 GHz frequency range with test field intensity amplitude of 3 V/m and 80% amplitude modulation with a 1 kHz sinus wave.

The test must be performed for at least one flow rate, with simulated input signals being allowed. No serious errors may occur during this test.

5.4.3.8    Resistance to interference caused by power lines, induced by high-frequency fields

Resistance to interference caused by power lines, induced by high-frequency fields is tested with the conversion apparatus switched on, between a frequency range of 150 kHz to 80 MHz at a test voltage of 3 V. Interference is applied to all inputs and outputs and to meter power.

The test must be performed for at least one flow rate, with simulated input signals being allowed. No serious errors may occur during this test.

6   Initial verification

6.1   In general

The initial verification process for an alcohol meter includes the following tests:

- visual inspection;

- functional tests.

6.2   Visual inspection

The following is assessed during a visual inspection of an alcohol meter:

whether the meter submitted for verification and its components correspond to the approved type, and whether they are not physically damaged,

- whether the meter and its components are satisfactory from the perspective of functionality and integrity:

the housing, especially at its seams, is whole and functional and does not bear signs of external interference with the meter;

the position and connection of counters corresponds to the approved type, bear no signs of external interference, and their function and recording exhibits no apparent difficulties;

the surface and seams of drums, ladles and all vessels exhibit no signs of leakage and porosity;

collection and insertion vessels and collection vessel taps exhibit no signs of leakage;

position marks are correctly set;

the meter has been installed according to manufacturer instructions.

6.3   Functional tests

6.3.1   Test equipment

Appropriate test equipment as per Section 5.3.1 is used.

6.3.2   Leak test

A drum leak test is performed for new and repaired meters according to Section 5.3.2.

If a leak is found, no further tests are performed on the meter.

6.3.3   Accuracy test

A meter accuracy test is performed according to Section 5.3.3.

The measurement error must not exceed the values listed in Table 1 for the relevant meter size.

6.3.4    Volume test for individual drum measurement chambers

A volume test for individual drum measurement chambers is performed according to Section 5.3.4.

The difference in volume between the largest and smallest chambers must meet the provisions of Section 2.3.2.1.

6.3.5   Resolution test

A resolution test is performed according to Section 5.3.5.

6.3.6    Backup measuring apparatus accuracy test

A resolution test is performed according to Section 5.3.6.

The measurement error for the backup drum must not exceed the values listed in Table 1 for the relevant meter size.

6.3.7    Structural element accuracy tests

Structural element accuracy tests are performed according to Section 5.3.7.

6.3.8   Display check

The check consists of:

checking that gearing used (after removing protective covers) is in accordance with meter documentation;

checking that the block preventing the display from being rolled back works properly.

6.4   Conversion apparatus test

If the alcohol meter has a built-in or auxiliary conversion apparatus, this conversion apparatus is tested from the perspective of its functionality and form the perspective of storage of alcoholometric table data in the conversion apparatus.

If the conversion apparatus receives input from connected measuring devices, such as a thermometer or densitometer for correction or conversion of measured data, they are subject to initial verification as part of the conversion apparatus.

6.5   Test assessment

If the meter passed all test requirements, official labels are placed on the meter on locations stipulated in the type approval certificate, and a verification sheet is issued for the meter.

7   Subsequent verification

7.1   In general

The subsequent verification process for an alcohol meter includes the following tests:

- visual inspection;

- functional tests.

7.2   Visual inspection

During subsequent verification, visual inspection is performed according to Section 6.2.

7.3   Functional tests

7.3.1   Test equipment

Appropriate test equipment according to Section 5.3.1 is used.

7.3.2   Leak test

A drum leak test is performed according to Section 5.3.2.

If a leak is found, no further tests are performed on the meter.

7.3.3   Accuracy test

A meter accuracy test is performed according to Section 5.3.3.

The measurement error must not exceed the values listed in Table 1 for the relevant meter size.

7.3.4   Display check

The check consists of checking that the block preventing the display from being rolled back works properly.

7.4   Conversion apparatus test

During subsequent verification, the conversion apparatus is tested according to Section 6.4.

7.5   Test assessment

If the meter passed all test requirements, official labels are placed on the meter on locations stipulated in the type approval certificate, and a verification sheet is issued for the meter.

8   Notified standards

To specify metrological and technical requirements for measuring devices and to specify verification test methods stemming from this measure of a general nature, the CMI notifies Czech technical standards, other technical standards or technical documents of international or foreign organizations or other technical documents containing detailed technical requirements (hereinafter "notified standards"). The CMI shall publish a list of these notified standards, with their assignation to relevant measures, along with the measure of a general nature, in a manner available to the public (on the website www.cmi.cz).

Compliance with notified standards or parts thereof is considered, to the extent and under the conditions stipulated by a measure of a general nature, to be compliance with those requirements stipulated by this measure to which these standards or parts thereof apply.

II.

J U S T I F I C A T I O N

The CMI issues, towards the implementation of Section 24c of Act No. 505/1990 Coll., on metrology, as amended, this measure of a general nature stipulating metrological and technical requirements for specified measuring devices and test methods for verification of these specified measuring devices.

In Decree No. 345/2002 Coll., specifying measuring devices for mandatory verification and measuring devices subject to type approval, as amended, the annex "List of Specified Measuring Device Types" classifies alcohol meters used to measure the amount of alcohol produced under item 1.3.8 as measuring devices subject to type approval and verification.

The CMI thus issues, towards the implementation of Section 24c of Act No. 505/1990 Coll., on metrology, as amended, for the measuring device type specified as "alcohol meters used to measure the volume of alcohol produced", this measure of a general nature, stipulating metrological and technical requirements for alcohol meters used to measure the volume of alcohol produced and test methods for type approval and verification of these specified measuring devices.

III.

E F F E C T I V E  D A T E

This measure of a general nature takes effect on the fifteenth day following the day of its publication (Section 24d of Act No. 505/1990 Coll., on metrology, as amended).

# Footnotes

1)   International Vocabulary of Metrology - Basic and General Concepts and Associated Terms (VIM)

2)   Act No. 61/1997 Coll., on alcohol, as amended

3) Ministry of Finance Notice No. 150/2008 Coll., on the supervision of the production and distribution of alcohol and on implementing other related provisions of the act on alcohol, Section 4(1)(c).

4)   OIML R 117-1 „Dynamic measuring systems for liquids other than water. Part 1: Metrological and technical requirements“ – available to the public at www.oiml.org

5)   OIML R 22 "International alcoholometric tables" – available to the public at www.oiml.org