1. ------IND- 2010 0349 CZ- EN- ------ 20100622 --- --- PROJET

The Czech Metrology Institute (CMI), as the competent local body for stipulating metrological and technical requirements for a specified measuring device and for the stipulation of test methods for type approval and for verification of a specified measuring device pursuant to Section 14(1) of Act No. 505/1990 Coll., on Metrology, as amended, and pursuant to provisions of Section 172 et seq. of Act No. 500/2004 Coll., the Administrative Code, as amended (hereinafter the "AC"), has on 26 February 2010 commenced ex officio proceedings pursuant to Section 46 of the AC, and based on supporting documents issues this:

I.

M E A S U R E  O F  A  G E N E R A L  N A T U R E

number: 0111-OOP-C009-10

reference no. 0313/002/10/Pos.,

publication date: 26 February 2010

stipulating metrological and technical requirements for specified measuring devices, including test methods for verifying specified measuring devices:

"automatic level gauges and automatic level gauges with automatic checking of metrological parameters"

1   Basic definitions

For the purposes of this measure of a general nature, terminology and definitions contained in the VIM1) and the following apply:

1.1   automatic level gauge, level gauge: an apparatus for the automatic measurement and display of fluid level with respect to a fixed reference point

1.2   automatic level gauge with automatic checking of metrological parameters: an automatic level gauge that is capable of automatically performing a regular check of metrological properties

1.3   fluid level sensor: an element that detects the surface of a fluid and provides information regarding its level

1.4   indicating device: a device that displays the measured fluid level value on an ongoing basis

1.5   transducer: a device that converts the fluid level value to fluid volume in a tank using the tank's calibration table

1.6   correction sensor: a sensor that measures relevant properties of the liquid and/or environment above the surface of the fluid in order to correct fluid level measurements

1.7   measurement range: a set of values stipulated by the manufacturer that can be measured by a given level gauge with a specified instrument uncertainty under defined conditions

1.8   dipping datum point: the intersection of the vertical measurement axis and the upper surface of the dip plate, or with the bottom surface of the tank if a dip plate is not provided. It constitutes the origin for the measurement of liquid levels

1.9   upper reference point: a clearly marked point above the surface of the measured liquid in the tank near the vertical measurement axis that passes through the base point, used as a reference to measure ullage

1.10   dip: the vertical distance between the dipping datum point and the liquid level in the tank

1.11   ullage: the vertical distance between the surface of the liquid in the tank and the upper reference point measured along the measurement axis

1.12   gauge reference length: the distance between the dipping datum point and the zero point of the level gauge

1.13   checking facility: a facility included in an automatic electronic level gauge that permits the following events to be detected and acted upon:

significant faults, and/or

incorrect functionality of a specific level gauge device, and/or

disruption of communication between individual level gauge devices

1.14   automatic checking facility: a checking facility that functions without the need for human intervention

1.15   permanent automatic checking facility (type P): An automatic checking facility that operates during each measurement cycle.

1.16   intermittent automatic checking facility (type I): An automatic checking facility that operates at certain time intervals or per fixed number of measurement cycles.

2   Metrological requirements

Metrological requirements are based on the requirements of recommendation OIML R 85 "Automatic level gauges for measuring the level of liquid in stationary storage tanks 2)".

2.1   Rated operating conditions

Automatic level gauges must function properly within an ambient temperature range specified by the manufacturer as an operating temperature range, which must be at least -25 °C to 55 °C

Level gauges must not be sensitive to relative humidity of the surrounding air, and must work properly at a relative humidity of up to 93%.

Reference laboratory conditions for level gauge tests are temperature (20 ± 5) °C and relative humidity of up to 85%.

2.2   Maximum permissible error

The maximum permissible error applies according to the automatic level gauge's measurement principle for indication of dip or ullage.

2.2.1    Maximum permissible error during verification

The maximum permissible measurement error of a level gauge before its installation in a tank is ± 1 mm.

2.2.2    Maximum permissible error during operation

The maximum permissible measurement error of a level gauge after its installation in a tank is ± 4 mm.

2.3   Hysteresis

Errors resulting from hysteresis during changes in surface level movement direction must not exceed 1 mm.

2.4   Resolution

The level gauge must have a resolution of less than 1 mm.

3   Technical requirements

Technical requirements are based on the requirements of recommendation OIML R 85 "Automatic level gauges for measuring the level of liquid in stationary storage tanks 2)".

3.1   Level gauge design

An automatic level gauge contains at least a fluid level sensor, a transducer and an indicating device.

It must be manufactured of solid, lasting materials without internal tensions, reasonably resistant to corrosion, the chemical effects of measured liquids and their vapours.

It must be capable of permanently resisting pressure, temperature and other operating conditions in the tank without functional failure or changes to metrological characteristics.

The minimum measurable distance to the fluid level from the dipping datum point must be known and indicated on the nameplate of the level gauge, if this value is greater than zero.

3.2   Fluid level sensor

For automatic level gauges with movable elements (e.g. float type), the movable element must be secured against intentional restriction of movement by the user. If such a movable element can be placed above or below the measured surface over a long term (e.g. for operational reasons), the level gauge must indicate that at that given instance it is not measuring the true surface level.

3.3   Indicating devices

If several indicating devices are used, each must meet the maximum allowable error requirements of Section 2.2. The difference between two indications under stable conditions must not be greater than 1 mm.

A remote indicating device must uniquely identify the corresponding automatic level gauge to which it belongs.

An automatic level gauge must display fluid level. The same display can show other measured values, such as ullage, as long as this data is automatically replaced by fluid level data within 10 seconds. Ullage data must be available for metrological purposes.

During regular operation, reading of measurement results must be reliable, easy and unambiguous.

Measurement results must contain the names or symbols of the units of length in which they are expressed. Indication device divisions must be 1 × 10n, 2 × 10n or 5 × 10n units of length, where n is a positive or negative whole number or zero.

3.4   Level gauge installation

Following installation, indication of measurement results must be easily accessible and legible.

The level gauge must be installed to permit checking of metrological parameter settings at the installation location.

If the tank has areas where the volume cannot be determined from dip and the tank's calibration table, the indicated dip values must be clearly differentiated from other values.

The level gauge must be installed so that the resultant influence of eddies, currents, turbulence, foam, condensation, changes to operating conditions, asymmetrical heating, wind and other influences on the level gauge's measurement is negligible. If needed, it must be protected from these influences.

An automatic level gauge must be installed on a tank in such a way that the discrepancy between the gauge reference length and the surface caused by movements of tank walls, bottom and top or stilling tubes does not exceed the maximum permissible operating error according to Section 2.2.2, otherwise these movements must be compensated for.

The level gauge must be installed in such a way that the fluid level sensor or measurement are not affected by obstructions in the tank

The effect of thermal expansion of the tank or structural tubing of the level gauge must be sufficiently small, or must be compensated for in such a way that temperature changes do not lead to the transgression of maximum permissible errors in measured fluid level.

If a correction sensor is available, it must be installed so that it measures appropriate values that influence magnitudes. If needed, several sensors may be installed to gain appropriate average correction values.

3.5   Auxiliary devices

Auxiliary devices that perform special functions in processing, transmitting or displaying measurement results may, for example, be connected to:

a printing device,

a memory device,

a conversion device.

Connection of such devices, be it temporary or permanent, must not change the metrological characteristics of the gauge. These devices must not have characteristics permitting fraudulent use of the gauge.

3.6    Automatic checking of metrological parameters

Automatic checking of metrological parameters consists of the periodic checking of metrological parameters of the level gauge under the following conditions:

during the check, the level gauge is compared to an independent length reference that is part of the measurement system,

it is performed automatically at least once a day or upon each movement of fluid in the tank, depending on which comes first,

the result is not used for level gauge measurement correction.

If the measurement error compared to the reference exceeds the maximum permissible error of the level gauge, the device must signal this state and record it in the gauge memory.

3.7   Ensuring measurement reliability

3.7.1   In general

An automatic level gauge must be equipped with a type I or P checking facility to prevent or signal significant errors.

3.7.2    Prevention or signalling of significant errors

An automatic level gauge must be designed and manufactured in such a way that it and its constituent parts do not exhibit significant errors when subjected to specified disruptive influences, or in such a way that these significant errors are detected by the checking facility and a visual and/or acoustical alarm is triggered that lasts until a user takes necessary steps to correct this error, or arranges for such steps to be taken.

During measurement, the checking facility must detect and immediately signal significant errors caused by the disruptive influences of:

radiated electromagnetic fields at radio frequencies,

conducted electromagnetic fields at radio frequencies,

electrostatic discharge,

bursts (transients) on signal, data and control lines,

surges on signal, data and control lines,

mains voltage dips, short interruptions and voltage variations,

bursts (transients) on mains

mains voltage ripple.

During measurement, the checking facility must detect and, after completing measurement, signal errors caused by the disruptive influences of:

variable humidity and temperature,

mains surges.

3.7.3    Signalling data loss or distortion

The checking facility must, by signalling the loss or distortion of data, permit the detection and addressing of:

improper functionality of the level gauge or a part thereof,

disrupted communication between level gauge devices or components.

If the checking facility detects the risk of data loss or distortion, an automatic visual and/or acoustical alarm is immediately triggered that lasts until a user takes necessary steps to correct this error, or arranges for such steps to be taken.

The design and construction of an automatic level gauge must ensure that data and commands permanently stored in the memory are always correct.

The integrity of all relevant measured data must be checked any time it is internally transmitted or stored, or sent via an interface to an external device, using measures such as:

parity bit,

check sum,

independent redundant storage,

a different communication protocol with repeated transmission.

3.8    Protection from unauthorized manipulation

An automatic level gauge must be built in such a way that the measuring or indicating apparatus cannot be interfered with in a way that could influence measurement integrity without visible damage to the verification mark.

4   Gauge labelling

4.1   Gauge labelling

The nameplate of an automatic level gauge must state the following:

the manufacturer's name or logo,

type,

serial number and year of manufacture,

type approval label,

measurement range

A remote indicating device must be labelled with the following:

type approval number,

tank identification.

4.2   Verification marks

The level gauge must have the necessary number of places for verification marks, allowing easy application and visual inspection without having to disassemble the device.

5   Type approval

5.1   In general

The approval process for an automatic level gauge includes the following tests:

external inspection;

laboratory tests of the level gauge;

tests of the resistance of the level gauge to disruptive influences of the surrounding environment;

a test of the automatic checking of metrological parameters function;

a test of the conversion unit.

5.2   External inspection

The following is assessed during an external inspection:

that prescribed technical documentation is complete;

that metrological and technical characteristics specified by the manufacturer in documentation conform to the requirements of this regulation listed in sections 2 and 3;

that the level gauge's functional wholes are complete and in accordance with prescribed technical documentation;

that the software version of the level gauge is the same as the version specified by the manufacturer.

5.3   Laboratory tests of the level gauge

5.3.1   Test equipment

The uncertainty of test equipment used to perform accuracy, resolution and hysteresis testing must be equal to or smaller than one fifth of the maximum permissible error according to Section 2.2.1.

5.3.2   Accuracy test

The accuracy of a level gauge is tested at least 10 fluid level points between the limits of the gauge measurement range, including its lower limit. At each point, measurement errors of the level gauge under test are calculated.

During the accuracy test, level gauge errors must not exceed the maximum permissible error listed in Section 2.2.1.

5.3.3   Resolution test

Level gauge resolution is tested at three fluid level points, including the lower limit of the level gauge range.

During the resolution test, level gauge errors must not exceed the maximum permissible error listed in Section 2.4.

5.3.4   Hysteresis test

The level gauge hysteresis test is performed at three fluid level points. During the hysteresis test, the fluid level is moved in both directions by at least 1/10 of the level gauge's measurement range.

During the hysteresis test, level gauge errors must not exceed the maximum permissible error listed in Section 2.3.

5.4   Tests of the resistance of the level gauge to disruptive influences of the surrounding environment

5.4.1    Tests of resistance to climactic influences

5.4.1.1    Limit temperatures test

Resistance to ambient temperature limits is tested with the level gauge switched on:

- with dry heat at a temperature of 55 °C for 2 hours,

- with cold at a temperature of -25 °C for 2 hours.

After the specified time has passed, an accuracy test at the given limit temperature is performed at three points of the measurement range, and hysteresis and resolution tests at one point of the measurement range of the level gauge. Measurement errors must not exceed the greatest permissible error according to Section 2.2.1, 2.3 and 2.4.

5.4.1.2   Humidity test

Resistance to humidity is tested on level gauges that have been switched off by cycling wet heat in two 24 hr cycles at a maximum temperature of 55 °C and a relative humidity greater than 95 %.

After both cycles have completed, no significant errors may occur, or they must be detected and corrected by the checking facility.

5.4.2   Limit power voltages test

Resistance to limit power voltage values are tested with the level gauge switched on. For DC, these limits are given by manufacturer specifications, for AC these limits are UN – 15 % and UN + 10 %, where UN is the nominal power voltage.

Measurement errors must not exceed the greatest permissible error according to Section 2.2.1 at limit power voltage values.

5.4.3   Electromagnetic compatibility (EMC) tests

5.4.3.1    Resistance to radiated electromagnetic fields at radio frequencies

Resistance to a radiated high-frequency electromagnetic field is tested with the level gauge switched on. The level gauge is subjected to an electromagnetic field in the 80 MHz to 2 GHz frequency range with test field intensity amplitude of 10 V/m and 80% amplitude modulation with a 1 kHz sinus wave.

During interference, at least one accuracy test must be performed at roughly halfway along the measurement interval. No serious errors may occur during this test, or they must be detected and corrected by the checking facility.

5.4.3.2   Resistance to interference caused by power lines, induced by high-frequency fields

Resistance to interference caused by power lines, induced by high-frequency fields is tested with the level gauge switched on, between a frequency range of 150 kHz to 80 MHz at a test voltage amplitude of 10 V. Interference is applied to all inputs and outputs and to gauge power.

During interference, at least one accuracy test must be performed at roughly halfway along the measurement interval. No serious errors may occur during this test, or they must be detected and corrected by the checking facility.

5.4.3.3   Resistance to electrostatic discharge

Resistance to electrostatic discharge is tested with the level gauge switched on, preferably with a contact discharge of 6 kV applied to a conductive section of the meter or with an 8 kV air discharge applied to coupling plates near the gauge.

During interference, at least one accuracy test must be performed at roughly halfway along the measurement interval. No serious errors may occur during this test, or they must be detected and corrected by the checking facility.

5.4.3.4    Resistance to bursts (transients) on signal, data and control lines

Resistance to rapid transient electrical phenomena/impulse groups is tested with the level gauge switched on at a test voltage of ± 1 kV on signal terminals with a repeat frequency of 5 kHz. Interference is applied to all gauge inputs and outputs.

During interference, at least one accuracy test must be performed at roughly halfway along the measurement interval. No serious errors may occur during this test, or they must be detected and corrected by the checking facility.

5.4.3.5    Resistance to surges on signal, data and control lines

Resistance to surges is tested with the level gauge switched on by applying at least three positive and three negative voltage surges:

asymmetrical line for line – line 0.5 kV,

asymmetrical line for line – ground 1.0 kV,

symmetrical line for line – ground 1.0 kV.

During interference, at least one accuracy test must be performed at roughly halfway along the measurement interval. No serious errors may occur during this test, or they must be detected and corrected by the checking facility.

5.4.3.6    Resistance to AC mains voltage dips, short interruptions and voltage variations

Resistance to short AC voltage gips, short interruptions and slow voltage variations are tested with the level gauge switched on at all AC power inputs by applying a voltage reduction:

of 100% of UN for 0.5 of an AC voltage cycle,

of 100% of UN for 1 AC voltage cycle,

of 60% of UN for 10 AC voltage cycles,

of 30% of UN for 25 AC voltage cycles,

of 20% of UN for 250 AC voltage cycles.

and a power outage simulated by a voltage reduction of 100% of UN for 250 AC voltage cycles, where UN is the nominal value of AC mains voltage.

During interference, at least one accuracy test must be performed at roughly halfway along the measurement interval. No serious errors may occur during this test, or they must be detected and corrected by the checking facility.

5.4.3.7   Resistance to bursts (transients) on mains

Resistance to rapid transient electrical phenomena/impulse groups is tested with the level gauge switched on at a test voltage of ± 2 kV on signal terminals for at least 1 minute for each amplitude and polarity with a repeat frequency of 5 kHz.

During interference, at least one accuracy test must be performed at roughly halfway along the measurement interval. No serious errors may occur during this test, or they must be detected and corrected by the checking facility.

5.4.3.8   Resistance to DC mains voltage dips, short interruptions and voltage variations

Resistance to short DC voltage gips, short interruptions and slow voltage variations are tested with the level gauge switched on at all DC power inputs. A series of three disturbances at least 10 seconds apart are applied, at the following voltage levels:

40 % of UN for 0.1 s;

70% of UN for 0.1 s;

0 % of UN for 0.01 s;

85 % of UN for u 10 s;

120 % of UN for 10 s.

During interference, at least one accuracy test must be performed at roughly halfway along the measurement interval. No serious errors may occur during this test, or they must be detected and corrected by the checking facility.

5.4.3.9    Resistance to DC voltage ripple

The voltage ripple is a sinusoid whose amplitude is 2% of mains voltage. The interference is applied for at least 10 minutes and for the entire time of the test of the gauge.

During interference, at least one accuracy test must be performed at roughly halfway along the measurement interval. No serious errors may occur during this test, or they must be detected and corrected by the checking facility.

5.4.3.10    Resistance to surges on mains

Interference is composed of three positive and three negative voltage impulses. For AC, pulses must be applied synchronously with AC power for phase shifts of 0°, 90°, 180° and 270°. Voltage amplitude is 1.0 kV for line – line and 2.0 kV for line – ground.

Immediately after interference, at least one accuracy test must be performed at roughly halfway along the measurement interval. No serious errors may occur during this test, or they must be detected and corrected by the checking facility.

5.5    Test of the automatic checking of metrological parameters

If the manufacturer declares that the level gauge is equipped with automatic checking of metrological parameters, it is investigated whether this function meets the requirements of Section 3.4. If this function meets prescribed parameters, this is stated in the type approval certificate, where the method of automatic checking of metrological parameters used is described, as well as the manner in which it should be tested during initial and subsequent verification.

5.6   Conversion unit test

If the level gauge is equipped with a built-in or auxiliary conversion unit that uses tank calibration tables to convert fluid level to fluid volume, also if it uses attached measuring devices for measuring temperature to automatically convert to volume at reference temperature, this conversion unit is tested from the perspective of its functionality and from the perspective of storage of tank calibration tables in the conversion unit.

If attached measuring devices (e.g. a thermometer or pressure meter) are used as inputs to the conversion unit for correction or conversion of measured data, they must be included in the type approval of the level gauge; if they are subject to type approval on their own, during level gauge type approval it must be assessed whether their connection does not influence the metrological characteristics of the level gauge.

6   Initial verification

6.1   In general

The following is performed during initial verification of level gauges:

a visual inspection of the level gauge and connected devices;

laboratory tests of the level gauge and connected devices;

installation in a tank, setting of the reference point, settings check;

sealing with a verification mark.

6.2   Visual inspection

During the visual inspection, the level gauge and connected devices submitted for verification are assessed as to whether they correspond to the approved type, and whether they are not physically damaged.

6.3   Laboratory tests of the level gauge

6.3.1   Test equipment

The uncertainty of test equipment used to perform accuracy, resolution and hysteresis testing must be equal to or smaller than one fifth of the maximum permissible error according to Section 2.2.1.

6.3.2   Accuracy test

The accuracy of a level gauge is tested at least 10 fluid level points between the limits of the gauge measurement range, including its lower limit. At each point, measurement errors of the level gauge under test are calculated.

During the accuracy test, level gauge errors must not exceed the maximum permissible error listed in Section 2.2.1.

6.3.3   Resolution test

Level gauge resolution is tested at three fluid level points, including the lower limit of the level gauge range.

During the resolution test, level gauge errors must not exceed the maximum permissible error listed in Section 2.4.

6.3.4   Hysteresis test

The level gauge hysteresis test is performed at three fluid level points. During the hysteresis test, the fluid level is moved in both directions by at least 1/10 of the level gauge's measurement range.

During the hysteresis test, level gauge errors must not exceed the maximum permissible error listed in Section 2.3.

6.3.5    Test of the automatic checking of metrological parameters

If the level gauge is equipped with automatic checking of metrological parameters, a test of this function is performed (according to the procedure listed in the type approval certificate).

6.3.6    Test of connected measuring devices

If attached measuring devices (e.g. a thermometer or pressure meter) are used as inputs to the conversion unit for correction or conversion of measured data, their verification is performed or arranged if their verification has not been performed independently and if the verification validity is not guaranteed for the validity period of the verification of the level gauge.

6.4    Setup and sealing at the installation location

After installation on the tank, the reference point is set, level gauge configuration and communication with connected measuring devices are checked, and the entire measuring system is sealed with verification marks.

7   Subsequent verification

Subsequent verification is performed using the same procedure as initial verification according to Section 6.

8   Notified standards

To specify metrological and technical requirements for measuring devices and to specify verification test methods stemming from this measure of a general nature, the CMI notifies Czech technical standards, other technical standards or technical documents of international or foreign organizations or other technical documents containing detailed technical requirements (hereinafter "notified standards"). The CMI shall publish a list of these notified standards, with their assignation to relevant measures, along with the measure of a general nature, in a manner available to the public (on the website www.cmi.cz).

Compliance with notified standards or parts thereof is considered to be, to the extent and under the conditions stipulated by a measure of a general nature, to be compliance with those requirements stipulated by this measure to which these standards or parts thereof apply.

II.

J U S T I F I C A T I O N

Towards the implementation of Section 24c of Act No. 505/1990 Coll., on metrology, as amended, the CMI issues this measure of a general nature, stipulating metrological and technical requirements for specified measuring devices and test methods for verification of these specified measuring devices.

In Decree No. 345/2002 Coll., specifying measuring devices for mandatory verification and measuring devices subject to type approval, as amended, the annex "List of Specified Measuring Device Types" classifies automatic level gauges and automatic level gauges with automatic checking of metrological parameters under item 1.1.7 as measuring devices subject to type approval and verification.

Towards the implementation of Section 24c of Act No. 505/1990 Coll., on metrology, as amended, for the measuring device type specified as "alcohol meters used to measure the volume of alcohol produced", the CMI thus issues this measure of a general nature, stipulating metrological and technical requirements for automatic level gauges and automatic level gauges with automatic checking of metrological parameters and test methods for type approval and verification of these specified measuring devices.

III.

E F F E C T I V E  D A T E

This measure of a general nature takes effect on the fifteenth day following the day of its publication (Section 24d of Act No. 505/1990 Coll., on metrology, as amended).

# Footnotes

1)   International Vocabulary of Metrology - Basic and General Concepts and Associated Terms (VIM).

2)   OIML R 85 "Automatic level gauges for measuring the level of liquid in fixed storage tanks" – available to the public at www.oiml.org