1. ------IND- 2010 0472 CZ- EN- ------ 20100907 --- --- PROJET

The Czech Metrology Institute (CMI), as the competent local body for stipulating metrological and technical requirements for a specified measuring device and for the stipulation of test methods for type approval and for verification of a specified measuring device pursuant to Section 14(1) of Act No. 505/1990 Coll., on Metrology, as amended, and pursuant to provisions of Section 172 et seq. of Act No. 500/2004 Coll., the Administrative Code, as amended (hereinafter the "AC"), has on 14 April 2010 commenced ex officio proceedings pursuant to Section 46 of the AC, and based on supporting documents issues this:

I.

M E A S U R E  O F  A  G E N E R A L  N A T U R E

number: 0111-OOP-C010-10

reference no. 0313/003/10/Pos.,

publication date: 21 May 2010

stipulating metrological and technical requirements for specified measuring devices, including test methods for verifying specified measuring devices:

"high-speed weigh-in-motion road vehicle scales"

1   Basic definitions

For the purposes of this measure of a general nature, terminology and definitions contained in the VIM1) and the following apply:

1.1   high-speed weigh-in-motion road vehicle scales, hereinafter "scales": automatic scales that measure the dynamic force on the tires of a moving vehicle and detect its presence on a pressure sensor over time, and calculate the total vehicle mass and load per axle or axle group, plus other vehicle parameters required by special legislation, while the vehicle is moving along its route

1.2   load receptor: a sensor of the dynamic force exerted by a vehicle tire on the road

1.3   weighing in entirety: determining the mass of a vehicle that is entirely on the load plate

1.4   static weighing: weighing the entire vehicle, load on axles or test load that is static without movement

1.5   weigh-in-motion: a process that uses measurement and analysis of dynamic forces on vehicle tires to determine the overall load exerted by a moving vehicle and the portions of this load that are transferred by the wheels or axles of this vehicle

1.6   vehicle mass: the entire mass of vehicle trailers including all elements permanently connected to the vehicle or located on it

1.7   axle: an axle containing two or more wheel assemblies spanning the entire width of the vehicle

1.8   axle group: a group of two or more axles on the same vehicle, defined by the total number of axles whose centre points are separated by less than a specified value

1.9   axle load: the portion of vehicle mass that is exerted via the axle on the load sensor during weighing

1.10   axle group load: total load on all axles that are part of an axle group

1.11   dynamic force on a tire: a force component changing over time applied perpendicular to the road surface by the tire on a moving vehicle; in addition to the force of gravity, this force can also include other dynamic effects on the moving vehicle

1.12   weighing range: the range between the minimum and maximum value of a measured variable magnitude, where scales function within the scope of given specifications

1.13   scale interval, d: the difference between two consecutive indicated or printed mass values during weigh-in-motion, expressed in units of mass

1.14   operating speed, v: the average speed of a moving vehicle as it passes over a load receptor, if it is to be weighed

1.14.1   maximum operating speed, vmax: the maximum vehicle speed for which a scale is designed for weigh-in-motion, and above which it is not guaranteed that the maximum permissible error will not be exceeded

1.14.2   minimum operating speed, vmin: the minimum vehicle speed for which a scale is designed for weigh-in-motion, and below which it is not guaranteed that the maximum permissible error will not be exceeded

1.14.3   operating speed range: the set of values between the minimum and maximum operating velocities for which the vehicle can be weighed in motion

1.14.4    maximum transit speed: the maximum speed that a vehicle can travel on the weigh zone without producing a shift in the performance characteristics of a weighing instrument beyond those specified.

1.15   maximum weighing limit (Max): the maximum weight the load sensor can measure during weigh-in-motion

1.16   minimum weighing limit (Min): the load value below which weigh-in-motion results can be subject to excessive relative error

1.17   reference scales: scales used for the static determination of the mass of a reference vehicle and the load on individual axles of the reference vehicle

1.18   vehicle: a road vehicle, under load or not, that is recognized by scales as a vehicle that is to be weighed

1.18.1   rigid vehicle: a road vehicle with a single chassis that includes neither coupling nor trailer, and that has two or more axles located along the length of the chassis

1.18.2   reference vehicle: a vehicle having a known conventional mass determined by reference scales

1.19    software subject to metrological verification of measuring devices: program(s), data and specific parameters of scale types that belong to a measuring device or equipment, and that define or perform functions that are subject to metrological verification of measuring devices

1.19.1   software parameter subject to metrological verification of measuring devices: a parameter of software of a measuring device or one of its modules that is subject to metrological verification of measuring devices

1.19.2    software identification: a sequence of legible characters that is a permanent part of the software (e.g. version number, checksum)

1.20   simulated functional test: a test performed on complete scales or parts thereof, where each part of the weighing operation is simulated

2   Metrological requirements

Metrological requirements are based on the requirements of recommendation OIML R 134-1 Automatic instruments for weighing road vehicles in motion and measuring axle loads - Part 1: Metrological and technical requirements – Tests2)

2.1   Service conditions

2.1.1   Operating temperature range

Scales must meet metrological requirements at temperatures from -20 °C to +40 °C.

2.1.2   Operating speed

Scales must meet applicable metrological requirements at vehicle velocities within the stipulated range of operating velocities. If the actual speed of a vehicle is outside the range of the scales' operating velocities during weighing, the scales must (see 3.4):

either automatically block measurement results from being issued, or

indicate or print out the value of the measured actual vehicle speed, and at the same time indicate or print out a clear warning that the measurement is outside the scales' operating speed range.

The operating speed must be indicated and/or printed out only after the entire vehicle has been weighed in motion.

The error in the indicated operating speed must not exceed 2 km/h.

2.2   Weighing range

The manufacturer of the scales specifies the weighing range, which is given by the values Max and Min, and must conform to the values listed in Table 1.

Table 1 – Weighing range

| Weighing range | Min kg | Max kg |
| --- | --- | --- |
| Axle load | 1000 | 20,000 |
| Vehicle mass | 3500 | ≥ 48,000 or as specified by special legislation |

2.3   Maximum permissible errors

2.3.1   Vehicle mass

The maximum permissible error for vehicle mass determined by weighing in motion is ± 5 %.

2.3.2    Axle load and axle group load

The maximum permissible error for axle and axle group load determined by weighing in motion is ± 11 %.

2.4   Units of measurement

The units of mass and load used by the equipment are the kilogram (kg) or tonne (t).

2.5   Scale interval

The scale interval must not exceed the values listed in Table 2.

Table 2 – Scale interval

| Axle load | 20 kg |
| --- | --- |
| Vehicle mass | 50 kg |

The scale interval of indication, recording or printing equipment must be in the form 1 × 10k, 2 × 10k or 5 × 10k, where k is a positive or negative whole number or zero.

2.6   Traffic intensity range

The manufacturer specifies the traffic intensity given by the maximum number of vehicles with mass above 3.5 t that can be weighed and recorded per hour and per day.

3   Technical requirements

Technical requirements are based on the requirements of recommendation OIML R 134-1 Automatic instruments for weighing road vehicles in motion and measuring axle loads - Part 1: Metrological and technical requirements – Tests2)

3.1   In general

Scales are an automatic measuring system composed of the following parts:

load receptors installed in the roadway,

vehicle recognition equipment,

equipment for measuring vehicle speed,

an indicator,

a printer,

a recording device,

equipment for the optical identification of vehicles,

auxiliary equipment,

that as a whole is capable of measuring dynamic forces on tires and detecting the presence of a moving vehicle on the load receptor over time, and of calculating total vehicle mass and axle or axle load values, speed, plus other vehicle parameters required by special legislation (e.g. axle separation, vehicle type).

The scales must be designed so that they are suitable for as great a range of vehicle types utilized for regular road use as possible.

The scales must be laid out and built in such a way that when they are correctly installed and used in  their intended environment, they maintain their metrological parameters to the extent specified by this legislation.

3.2   Load receptor

The load receptor installed in the roadway must detect the presence of a moving vehicle and record the dynamic forces on the tires.

3.3   Vehicle recognition equipment

Scales are intended for operator-free operation, and thus must have vehicle recognition equipment. This equipment must detect the presence of a vehicle in the weighing zone and must ascertain when the entire vehicle was weighed. The equipment must permit the categorization of vehicles in accordance with special transport legislation 3).

The scales must not indicate, record or print out the mass of the vehicle if all of the vehicle's wheels were not weighed or the vehicle was not recognized.

3.4   Equipment for measuring vehicle speed

The speed of a vehicle during weighing must be indicated, and if appropriate recorded and printed out as part of the vehicle weighing record, in km/h, after being rounded to the nearest whole number.

The scales must not indicate, record or print out the mass or axle or axle group values for any vehicle that has passed over the load receptor at a speed outside the specified range of operating speeds without an accompanying clear warning signal that these results are not verified.

3.5   Indication equipment

3.5.1   Indication during normal operation

Displayed mass and load values must be indicated precisely and unambiguously, and must not lead to errors.

Scales are intended for indication of the following values:

measured values of total mass, including the unit of measurement,

measured values of axle or axle group load, including the unit of measurement,

maximum permissible total mass value, including the unit of measurement,

maximum permissible axle or axle group load, including the unit of measurement,

the speed of the vehicle,

the time (the date, hour, minute and second the measurement was performed), if appropriate.

If scales in a corresponding application are not intended and approved for total mass or axle or axle group load, the scales can indicate this value for information only, clearly marked that this indication is not in the verified range of measurement.

3.5.2    Limit indication values

Scales may indicate, record or print out data on measurement of axle load, axle group load or vehicle mass, if single-axle load (partial weighing) is lower than Min or greater than Max only if the indication and/or printed output is accompanied by a clear warning of this problem.

3.6   Printing device

3.6.1    Print output during normal operation

The printed results must be correct, suitably identified and unambiguous. Printing must be clear, legible, indelible and permanent.

The contents of the printed output ensues from the values for which the scales are intended, as follows:

measured values of total mass, including the unit of measurement,

measured values of axle or axle group load, including the unit of measurement,

maximum permissible total mass value, including the unit of measurement,

maximum permissible axle or axle group load, including the unit of measurement,

the time (resolution to the second) and date (day, month, year),

the speed of the weighed vehicle,

the type of scales (e.g. an abbreviation).

If scales in a corresponding application are not intended and approved for total mass or axle or axle group load, the scales can indicate this value for information only, clearly marked that this indication is not in the verified range of measurement.

3.6.2    Agreement between indication and printing devices

For the same load, there must not be any difference between the indicated and printed weighing result, if the indication and printing devices have the same scale interval.

3.7   Recording device

3.7.1    Scope of recorded values

The recording device must record and maintain all store measurement data. The minimum scope of data recorded and stored by this device is the same as the contents of printed output per 3.6.1.

3.7.2   Data storage

Data may be stored in the scales' memory (e.g. on a hard disk) or in external storage for subsequent operations (indication, printing, transmission, summation, etc.). Stored data must be reasonably protected against intentional and unintentional changes during the transmission and/or storage process, and must contain all relevant information needed to reconstruct past measurements.

The following stored data security requirements apply:

relevant security requirements in 3.14,

the software transmission and download process must be secured in accordance with requirements in 3.14,

external memory identification and security attributes must ensure integrity and authenticity,

exchangeable storage media for storing measurement data need not be sealed, providing that stored data is secured by a specific checksum or key code,

when storage capacity is exhausted, overwriting data with new data must be possible only using a key code that is available only to authorized individuals.

3.8    Optical vehicle recognition device

Scales must be equipped with a device according to 3.8.1 or 3.8.2 for the unambiguous recognition of those vehicles that during weighing were found to exceed specified mass parameters. This recognition must meet security, integrity and authenticity requirements.

3.8.1   Imaging unit

Scales may be equipped with a digital camera that captures the weighing situation with safe identification of the weighed vehicle, which is displayed on a connected remote display along with the following values measured by the scales:

measured values of total mass, including the unit of measurement,

measured values of axle or axle group load, including the unit of measurement,

optionally also:

maximum permissible total mass value, including the unit of measurement,

maximum permissible axle or axle group load, including the unit of measurement.

3.8.2    Image recording unit

Scales may be equipped with an image recording unit that must capture the situation during weighing, ensuring identification of the vehicle being weighed.

An image-recording unit working in automatic mode should enable the setting of a limit mass for image recording.

The situation on the scales is recorded by a digital camera, which outputs individual digital images or video sequences stored in digital memory.

The following must be displayed on individual images or in video sequences in the data display field:

measured values of total mass, including the unit of measurement,

measured values of axle or axle group load, including the unit of measurement,

maximum permissible total mass value, including the unit of measurement,

maximum permissible axle or axle group load, including the unit of measurement,

the time (resolution to the second) and date (day, month, year),

the speed of the weighed vehicle,

the type of scales (e.g. an abbreviation).

For digital images, image information and information regarding measured values must be inseparably joined into one data file and integrated into the pixel structure of the digital image. To ensure integrity, the digital image data file must have a digital mark (signature). The origin (authenticity) of the entire digital image data file must be uniquely identifiable via coding (e.g. the ID number of the scales).

The integrity and authenticity of video sequences that are to be archived must be ensured to prevent undue changes to the content of images and measured data, or improper assignation.

3.9   Auxiliary devices

Any external devices connected to the scales via an appropriate interface must not have a negative influence on their metrological parameters.

3.10    Resistance to external influences

Disruptive external influences on the scales must not lead to measurement errors that would exceed the scales' greatest permissible error according to 2.3.1 or 2.3.2.

3.10.1   Physical robustness

The design of the scales and materials used must guarantee sufficient rigidity, stability and resistance to mechanical vibrations and shocks. The manufacturer must specify the physical conditions in which the scales should be used. For the load receptor, a physical environment class with high or very high level of vibrations and shocks applies, designated as M3 pursuant to Government Decree No. 464/2005 Coll., stipulating technical requirements for measuring instruments4).

3.10.2   Weather resistance

In its off state, a load receptor installed in a roadway must be able to withstand temperatures between -40 °C and 70 °C without damage, and after returning to its operating temperature range, must function within the range of maximum permissible error.

In order to ensure correct measurement in relation to ambient temperature and the scales' operating temperature range, the scales must have a temperature measurement device. The scales must be capable of automatically recognizing a temperature outside the operating temperature range, and displaying a suitable warning. Any weighing taking place at that moment must be terminated and the scales must block further weighing or switch themselves off.

The scales must not be sensitive to the relative humidity of the surrounding air.

3.10.3    Dust and water resistance

Those portions of the scales that are subject to the effects of weather must have at least an IP 67 housing to provide protection from dust and temporary immersion in water, and other parts at least IP 54.

3.10.4   Electromagnetic compatibility (EMC)

The scales must not be influenced by electrical or electromagnetic interference, or must react to it in a defined manner (e.g. reporting an error, blocking measurement, etc.). They must also not radiate any unwanted electromagnetic fields.

During laboratory EMC tests, the scales or parts thereof must exhibit normal function and the results of simulated functional tests must be within the limits of greatest permissible error according to 2.3.1 or 2.3.2.

3.11   Power

Scales powered from the electrical grid must meet metrological requirements under normal voltage fluctuation conditions.

If voltage declines below the minimum operating voltage, the scales must be blocked from working or their activity outside specified operating conditions must be clearly indicated, for example by a suitable warning.

3.12    Measuring device security and fraud prevention

Scales must not have characteristics that would facilitate fraudulent use, and there must be a minimum of ways in which they can be unintentionally improperly used. Components that are not intended to be disassembled or adjusted by the user must be protected from such activity.

3.13   Software

3.13.1   Software subject to metrological verification of measuring devices

This software used in scales must be presented in such a form that software cannot be changed without damaging a seal, or each change in software can be automatically recorded and its nature specified with the use of an ID code.

Documentation for software for scales must contain:

a description of the software subject to metrological verification of measuring devices,

a description of measurement algorithm accuracy (e.g. programming modes),

a description of the user interface, menus and dialogs,

unique identification of the software,

a description of included software (e.g. operating environment),

a hardware system overview, e.g. a topological block diagram, type of computer(s), source code for software functions, etc., if not described in the user manual,

software security resources

a user manual.

3.13.2   Software security resources

Resources for securing software subject to metrological verification of measuring devices are as follows:

- only authorized individuals may be given access, for example using codes (passwords) or a special device (hardware key, etc.); codes must be changeable,

- the measuring device's memory must store all accesses, listing the date of the access, identification of the authorized individual performing the access, and the type of access,

- memory capacity must be sufficient for at least 2 years of expected accesses; if memory capacity for access record storage is exhausted, no automatic erasure of any stored records can take place,

- it must be possible to recall relevant access records to the full extent of information recorded,

- it must not be possible to erase access records without removing a physical seal,

- downloading of software subject to metrological verification must be possible only via an appropriate secure interface connected to the scales,

- the software must include identification of its version, which must change if any software version changes occur,

- functions that are performed or launched via a software interface must meet the terms and conditions of this legislation.

3.14    Hardware and software security

3.14.1   In general

All scale equipment including software, which is intentionally to be prevented from disconnection or removal by a user or other individual, must be equipped with a housing or other suitable security means. It must be possible to seal housings after their closure; sealing points must be easy to access in all instances. All parts of the measuring system that cannot be protected by housings must be equipped with sufficiently effective means of preventing operations that tend to influence measuring accuracy.

Each piece of scale equipment that could influence measuring results, especially equipment for  calibration and adjustment of scales or for correction of measured values, must be sealed.

3.14.2    Means of security

Security must consist of sealed housings, encryption, passwords or similar software means in such as way that:

- the software security requirements in 3.13.2 apply,

- transmission of measurement results data via an interface must be protected from intentional, unintentional and random changes,

- the scales must be secured in such a way to make it possible to secure scale settings separately,

- stored data must be protected from intentional, unintentional and random changes.

3.15   Installation

3.15.1   In general

Scales must be installed so as to minimize any adverse effects of the installation environment on measuring accuracy and related data. The weighing area must be kept free of all substances that could affect the accuracy of the scales. All installation requirements that have an effect on the weighing operation (e.g. site levels, roadway geometry) must be stipulated in sufficient detail.

3.15.2   Roadway geometry

The section of roadway 50 m in front of and 25 m past the load receptor must meet the following requirements:

- the road gradient must be ( 1% and constant,

- the road's cross gradient must be ( 3%,

- the radius of curvature of the road's longitudinal axis must be ≥ 1000 m,

- the road must be free of irregularities causing local gradient changes,

- the depth of wear troughs must not be greater than 4 mm.

Scales must not be installed in areas where frequent acceleration or slowing could occur, and must not be installed in areas where the number of traffic lanes changes.

3.15.3   Road characteristics

The road in the load receptor installation location must meet the following requirements:

- below the road's top layer, there must be a standard substrate without reinforced locations (e.g. without special facilities such as service shafts, etc.),

- load receptors must be installed in homogenous layers where the surface is undamaged,

- along the entire span of load sensors, the road must be homogenous across each traffic lane and without masonry joints,

- load receptors must not be installed in locations where undesirable dynamic effects could occur, such as on bridges.

4   Markings

4.1   Markings on the measuring device

Scales must bear the following markings:

identification mark of the manufacturer,

type designation,

serial number,

if scales are not suited or intended for weighing vehicles that have specific characteristics (e.g. axle suspension design, number of axles) or transporting loads with specific characteristics (e.g. liquids), they must be marked with a notice or this restriction on suitability for weighing with a clear specification of the type and scope of this restriction (if relevant for the scales in question),

maximum transit speed, in km/h,

direction of weighing (if applicable),

power supply voltage, in V,

power supply frequency, in Hz,

temperature range (when not –20 °C to +40 °C), in °C,

software identification (if applicable),

and the following information regarding metrological parameters:

upper weighing limit Max = ....., in kg or t;

lower weighing limit Min = ....., in kg or t;

scale interval d = ....., in kg or t;

maximum operating speed vmax = ....., in km/h;

minimum operating speed vmin = ....., in km/h;

maximum number of axles on a vehicle (if applicable) Amax;

type approval mark in accordance with national requirements.

4.2   Presentation of markings

Markings according to 4.1 must be legible, unambiguous, comprehensible and indelible under normal usage conditions.

These markings may be either in the Czech language or in the form of adequate, internationally agreed and published pictograms or signs.

Markings shall be grouped together in a clearly visible place on the scales, either on a descriptive plate or sticker fixed permanently near the indicating device, or on a non-removable part of the scales themselves. In case of a plate or sticker which is not destroyed when removed, a means of

securing shall be provided, e.g. a non-removable control mark that can be applied.

4.4   Control markings

Scales and components thereof must permit the placement of a control marking or markings so that:

scale components on which markings are placed cannot be removed from the scales without damaging the markings,

the markings can be placed without changing the metrological characteristics of the scales,

they are visible on the scales during normal installation.

5   Measuring device type approval

The following are performed during type approval:

external inspection,

tests of the scales' resistance to disruptive environmental effects,

functional weigh-in-motion tests on location during road traffic.

5.1   External inspection

During external inspection, the following is assessed:

- that prescribed technical documentation is complete, including a user manual,

- that metrological and technical characteristics specified by the manufacturer in documentation comply with the requirements of this legislation specified in clauses 2 and 3,

- that functional wholes are complete and comply with prescribed technical documentation,

- that the scales' software version agrees with the version specified by the manufacturer.

5.2    Simulated functional tests in the laboratory

Simulated functional tests are performed when assessing resistance to the influence of the external environment according to 5.3, on complete scales unless the size and/or configuration of the scales makes it impossible to test them in their complete form. In such cases, testing is allowed with a load signal generator taking the place of load receptors.

The metrological body approving measuring device types can accept a manufacturer's proposal to modify the method and manner in which simulated functional tests are performed, if suitable with regards to the specifics of the technology and design of the scales' measurement chain.

5.3   Tests of resistance to the influence of the external environment

5.3.1    Tests of resistance to physical effects

5.3.1.1    Tests of resistance to random vibrations

Resistance to random physical vibrations is tested on scales in their on state by applying vibrations with the following parameters:

frequency range: 10 Hz to 150 Hz,

overall effective acceleration level: 7 m/s2,

spectral density of acceleration 10 Hz to 20 Hz: 1 m2/s3,

spectral density of acceleration 20 Hz to 150 Hz: -3 dB/octave,

in all three axes, always for a period of 2 minutes.

During the course of this test the tested scales must remain functional, and during the following simulated functional test, measurement error must not exceed the maximum permissible error listed in 2.3.1 or 2.3.2.

5.3.1.2   Shock resistance test

Resistance to shocks is tested with the scales switched on, by applying repeated shocks with the following parameters:

peak acceleration: 100 m/s2,

nominal impulse duration: 16 ms,

corresponding velocity change: 1 m/s,

number of shocks in each direction: 1000 ± 10.

During the course of this test the tested scales must remain functional, and during the following simulated functional test, measurement error must not exceed the maximum permissible error listed in 2.3.1 or 2.3.2.

5.3.2   Weather resistance tests

5.3.2.1    Resistance to limit temperatures

Resistance to limit ambient temperatures according to 3.10.2 is tested with the scales switched off:

- with dry heat at 70 °C for 2 h,

- with cold at -40 °C for 2 h.

Following this test, the scales must not exhibit any damage, and during the following simulated functional test, measurement error must not exceed the maximum permissible error listed in 2.3.1 or 2.3.2.

5.3.2.2    Resistance to operating temperatures

Resistance to operating ambient temperatures according to 2.1.1 is tested with the scales switched on:

- with dry heat at the upper limit of the operating temperature range for 2 h,

- with cold at the lower limit of the operating temperature range for 2 h.

During this test the scales must function normally, and the error during a simulated functional test must not exceed the maximum permissible error listed in 2.3.1 or 2.3.2.

5.3.2.3   Resistance to air humidity

Resistance to air humidity according to 2.2 is tested with the scales switched on with two 24-hour cycles of moist heat with a maximum temperature of 40 °C.

During this test the scales must function normally, and the error during a simulated functional test must not exceed the maximum permissible error listed in 2.3.1 or 2.3.2.

5.3.2.4    Dust and water resistance

Dust and water resistance according to 3.10.3 is tested in the off state on those parts of the scales that are exposed to the effects of weather.

Following this test, the scales must not exhibit any damage, and during the following simulated functional test, measurement error must not exceed the maximum permissible error listed in 2.3.1 or 2.3.2.

5.3.3   Electromagnetic compatibility (EMC) tests

5.3.3.1   Resistance to interference caused by power lines, induced by high-frequency fields

Resistance to interference caused by power lines, induced by high-frequency fields is tested with the scales switched on, over a 150 kHz to 80 MHz frequency range with a test field amplitude of 10 V. The interference is applied to signal cabling longer than 3 m, at all DC power inputs and outputs, at all AC power inputs and outputs, and at all functional ground connections.

During a simulated functional test under the given test conditions, measurement error must not exceed the maximum permissible error listed in 2.3.1 or 2.3.2., or the system must detect a serious error and react to it.

5.3.3.2   Resistance to radiated high-frequency electromagnetic fields at radio frequencies

Resistance to a radiated high-frequency electromagnetic field is tested with the scales switched on, in the 80 MHz to 2 GHz frequency range with test field intensity amplitude of 10 V/m and 80% amplitude modulation with a 1 kHz sinus wave. The interference is applied to all sides of the scales' housing.

During a simulated functional test under the given test conditions, measurement error must not exceed the maximum permissible error listed in 2.3.1 or 2.3.2., or the system must detect a serious error and react to it.

5.3.3.3  Resistance to electrostatic discharge

Resistance to electrostatic discharge is tested with the scales switched on, preferably with a contact discharge of 6 kV or with an 8 kV air discharge. Discharges are applied to the scales' housing and to coupling plates near the scales.

During a simulated functional test under the given test conditions, measurement error must not exceed the maximum permissible error listed in 2.3.1 or 2.3.2., or the system must detect a serious error and react to it.

5.3.3.4   Resistance to electrical fast transients/bursts

Resistance to electrical fast transients/bursts is tested with the scales switched on at a no-load test voltage of ± 1 kV for at least 1 minute on signal terminals with a repeat frequency of 5 kHz. The interference is applied to signal cabling longer than 3 m, at all AC power inputs and outputs, and at all functional ground connections longer than 3 m.

During a simulated functional test under the given test conditions, measurement error must not exceed the maximum permissible error listed in 2.3.1 or 2.3.2., or the system must detect a serious error and react to it.

5.3.3.5    Resistance to electrical surges

Resistance to surges is tested with the scales switched on by applying a surge of:

± 1 kV line to ground; to signal lines longer than 30 m;

± 0.5 kV line to line and a symmetrical test voltage of ± 0,5 kV on DC power lines longer than 10 m.

Interference is applied to signal lines longer than 30 m or to lines partly or entirely installed outdoors, regardless of their length.

During a simulated functional test under the given test conditions, measurement error must not exceed the maximum permissible error listed in 2.3.1 or 2.3.2., or the system must detect a serious error and react to it.

5.3.3.6    Resistance to power-frequency magnetic fields

Resistance to 50 Hz magnetic fields is tested with the scales switched on with a continuous magnetic field of intensity 30 A/m on the equipment housing.

During a simulated functional test under the given test conditions, measurement error must not exceed the maximum permissible error listed in 2.3.1 or 2.3.2., or the system must detect a serious error and react to it.

5.3.3.7   Resistance to AC mains voltage dips

Resistance to short AC mains voltage dips, short interruptions and slow voltage variations is tested with the scales switched on at all AC power inputs with a supply current of < 16 A by applying a voltage reduction:

of 40% of UN for 10 AC voltage cycles,

of 70% of UN for 25 AC voltage cycles,

of 80% of UN for 250 AC voltage cycles.

where UN is the nominal value of AC mains voltage.

During a simulated functional test under the given test conditions, measurement error must not exceed the maximum permissible error listed in 2.3.1 or 2.3.2., or the system must detect a serious error and react to it.

5.3.4    Limit power voltages test

Resistance to limit power voltage values is tested with electronic devices switched on. For AC these limits are UN – 15 % and UN + 10 %, where UN is the nominal power voltage.

During a simulated functional test under the given test conditions, measurement error must not exceed the maximum permissible error listed in 2.3.1 or 2.3.2., or the system must detect a serious error and react to it.

5.4    Weigh-in-motion accuracy road tests

5.4.1   In general

Complete scales installed at a weighing location agreed upon with the client are subjected to weigh-in-motion accuracy tests.

5.4.2   Test equipment

5.4.2.1   Reference vehicles

Reference vehicles shall be selected for weigh-in-motion tests so that they represent various axle configurations, truck/trailer configurations, truck/trailer connection systems and suspension systems.

The following reference vehicles are used:

a rigid two-axle vehicle,

and at least two other reference vehicles that must be selected from the three options listed below:

one rigid three-axle or four-axle vehicle,

one vehicle with four or more connected axles,

one two-axle or three-axle rigid vehicle and a two-axle or three-axle drawbar trailer.

As an alternative to the preceding provisions of this clause, vehicles directly from the stream of traffic can be used for test purposes described in 5.4.6, as long as they are subsequently subjected to static weighing to determine the reference mass value.  In this case, a set of 100 or more vehicles is considered to sufficiently represent various axle configurations.

5.4.2.2    Reference vehicle loading

During testing, reference vehicles must be used both unloaded, and loaded in a way that two significantly different load levels are used.

If the scales are to be used to determine vehicle mass or single-axle or axle group loads for vehicles transporting loads whose centre of gravity can shift during vehicle motion, reference vehicles must include those carrying loads in the form of liquids or other products that can change the position of their centre of gravity during vehicle motion.

5.4.2.3   Reference scales

During testing, standalone reference scales must be available to determine the conventional true mass value of each vehicle and the reference single-axle or axle group load.

5.4.2.3.1    Reference scales for measuring reference vehicle mass

To determine the conventional overall mass value of reference vehicles, preference is given to stand-alone reference scales that can determine the conventional mass value of each reference vehicle by weighing it all at once with an error less than or equal to one third of the applicable MPE for weigh-in-motion in accordance with 2.3.1.

5.4.2.3.2    Reference scales for measuring reference axle load of a reference vehicle

To determine the conventional axle load, stand-alone portable reference scales for weighing Class III or IV vehicles are used, or low-speed scales of accuracy class 1 or better.

5.4.3   Test conditions

Tests are performed under operating conditions stipulated by the manufacturer of the scales in question.

5.4.4    Determining reference values for reference vehicles

5.4.4.1   Conventional true value of the reference vehicle mass

The conventional true mass value for each reference vehicle, unloaded and loaded, must be determined by static full-draught weighing on reference scales (see 5.4.2.3.1).

5.4.4.2    Conventional true value of the axle load of a reference vehicle

Reference scales are used to determine conventional axle load values for an unloaded and loaded reference vehicle (see 5.4.2.3.2).

The reference scales are used to sequentially determine the load on each axle of the reference vehicle, with at least three to five test runs in both directions. The mean reference axle load is calculated as the arithmetic average of recorded values.

To correct for the influence of the method used, the total vehicle mass is calculated by axles as the sum of mean load values on individual axles VM.

The corrected mean reference load value per axle is then:

(1)

where: VMref 	is	the conventional value of each reference vehicle mass determined by full-draught 			weighing in accordance with 6.2.2.1.

To verify that reference axle loads are correct, the following must apply:

The corrected mean load value (see above) is used as the conventional per-axle load of the reference vehicle.

5.4.5    Verification of scales installation at the weighing location

Road geometry is checked according to 3.15.2, and must meet all criteria.

5.4.6    Weigh-in-motion reference vehicle tests

5.4.6.1   Test runs

Each reference vehicle, unloaded and then loaded, must perform at least five test runs at each of the three following speeds:

near the maximum operating speed, vmax,

near the minimum operating speed, vmin,

near the middle of the operating speed range,

(each reference vehicle must thus perform a total of 30 test runs).

For every five test runs at a given tests speed, the vehicle must be positioned above the centre of the load receptor three times, once on the left and once on the right side of the load receptor.

When using vehicles directly from the stream of traffic in accordance with 5.4.2.1, the number of test runs equals the number of vehicles used, thus at least 100.

5.4.6.2   Test run speed

Vehicle speed must be kept as constant as possible during each test run. Scales must indicate and record the speed of the tested vehicle as it passes over the load receptors.

5.4.6.3   Weigh-in-motion accuracy tests

During weigh-in-motion accuracy tests, all test runs are performed according to 5.4.6.1 using reference vehicles in accordance with 5.4.2.1. The values of all vehicle mass indications and all axle load indications are recorded. For each recorded value (total vehicle mass, axle or axle group load), the relative error d is calculated in percent:

(2)

where:	C is the value measured by the scales,

R is the corresponding reference value measured by the reference scales.

The number of relative errors d that exceed the stipulated maximum permissible error according to 2.3.1 or 2.3.2 for each quantity is determined, and this number is expressed as the relative number of values for each quantity as follows:

(3)

where:	n is the number of calculated differences exceeding the maximum permissible error,

N is the total number of recorded values for the given quantity.

The number of relative errors exceeding the maximum permissible error Pde must not be greater than 5%, with their distribution among individual vehicle types being recorded.

5.4.7    Operating speed tests

5.4.7.1    Operating speed blocking test

During the operating speed blocking test, a test run by one reference vehicle must take place at a speed outside the operating speed range, as follows:

- at a speed at least 5% higher than the maximum operating speed, vmax,

- at a speed at least 5% lower than the minimum operating speed, vmin, (if the scales can be used for this).

The scales must detect the above conditions and must react in accordance with 3.4.

5.4.7.2    Operating speed test

To determine and test operating speed during a weigh-in-motion test, six test runs shall take place with an unloaded two-axle rigid reference vehicle across load receptors at a constant speed. Three runs must take place near the maximum operating speed vmax, and three additional runs must take place at exactly the listed minimum operating speed vmin.

The reference speed value to be used to calculate the error of the indicated operating speed for each test run must be proportional to the distance between the measured axles (to the nearest 10 mm) of the stationary two-axle rigid reference vehicle divided by the measured time interval (to the nearest millisecond) between arrival at a defined point (e.g. the first load receptor) and the rear axle of the moving two-axle rigid reference vehicle.

The indicated operating speed error must not exceed the error given in 2.1.2.

5.4.8    Vehicle recognition device blocking test

Two reference vehicles are connected by a suitable towing device that ensures constant distance in order to create an atypical sequence of trailers between the rear axle of the pulling vehicle and the front axle of the pulled vehicle.

The scales must then:

recognize the combined vehicle as two vehicles,

determine the correct axle group load or detect an error,

detect the fact that all weighed wheels of the vehicle were not weighed during the weighing operation, and not indicate or print any mass or load values if the indication and/or printout has a clear warning.

6   Initial verification

6.1   In general

The following is performed during initial verification:

a visual inspection,

functional weigh-in-motion tests in road traffic,

operating speed tests,

a vehicle recognition device blocking test

6.2   Visual inspection

The following is assessed during a visual inspection of scales submitted for verification:

conformance of the scales with the approved type,

completeness and condition of the scales' functional wholes,

that the software version an approved one.

6.3   Functional weigh-in-motion tests in road traffic

During initial verification, a scales accuracy test is performed with by weighing in motion in road traffic at the scales installation location per 5.4.6, with the test scope being restricted to two operating speeds. The following reference vehicles are used:

one rigid two-axle vehicle,

and at least two other reference vehicles that must be selected from the three options listed below:

one rigid three-axle or four-axle vehicle,

one vehicle with four or more connected axles,

one two-axle or three-axle rigid vehicle and a two-axle or three-axle drawbar trailer.

Alternately, this test can be performed using vehicles directly from the stream of traffic as reference vehicles; in this case, the number of such vehicles must be at least 60.

6.4    Operating speed tests

During initial verification, operating speed tests in accordance with 5.4.7.1 are performed.

6.5   Vehicle recognition device blocking test

During initial verification, vehicle recognition device blocking tests in accordance with 5.4.8 are performed.

7   Subsequent verification

7.1   In general

The following is performed during subsequent verification:

- a visual inspection,

- functional weigh-in-motion tests in road traffic.

7.2   Visual inspection

Visual inspection during subsequent verification proceeds in accordance with 6.2.

7.3   Functional weigh-in-motion tests in road traffic

During this verification, a scales accuracy test is performed with by weighing in motion in road traffic at the scales installation location per 6.3 is performed, with the test scope being restricted to one reference vehicle load used, which is the greater of the two loads used in accordance with 5.4.2.2, with the total number of runs having to be at least 60.

Alternately, this test can be performed using vehicles directly from the stream of traffic as reference vehicles in accordance with 5.4.6.1; in this case, the number of such vehicles must be at least 60.

8   Notified standards

To specify metrological and technical requirements for measuring devices and to specify verification test methods stemming from this measure of a general nature, the CMI notifies Czech technical standards, other technical standards or technical documents of international or foreign organizations or other technical documents containing detailed technical requirements (hereinafter "notified standards"). The CMI shall publish a list of these notified standards, with their assignation to relevant measures, along with the measure of a general nature, in a manner available to the public (on the website www.cmi.cz).

Compliance with notified standards or parts thereof is considered to be, to the extent and under the conditions stipulated by a measure of a general nature, to be compliance with those requirements stipulated by this measure to which these standards or parts thereof apply.

II.

J U S T I F I C A T I O N

The CMI issues, towards the implementation of Section 24c of Act No. 505/1990 Coll., on metrology, as amended, this measure of a general nature, stipulating metrological and technical requirements for specified measuring devices and test methods for verification of these specified measuring devices.

In Decree No. 345/2002 Coll., specifying measuring devices for mandatory verification and measuring devices subject to type approval, as amended, the annex "List of Specified Measuring Device Types" classifies high-speed weigh-in-motion road vehicle scales under item 2.xx.xx as measuring devices subject to type approval and verification.

The CMI thus issues, towards the implementation of Section 24c of Act No. 505/1990 Coll., on metrology, as amended, for the measuring device type specified as "high speed weigh-in-motion road vehicle scales ", this measure of a general nature, stipulating metrological and technical requirements for high speed weigh-in-motion road vehicle scales and test methods for type approval and verification of these specified measuring devices.

III.

E F F E C T I V E  D A T E

This measure of a general nature takes effect on the fifteenth day following the day of its publication (Section 24d of Act No. 505/1990 Coll., on metrology, as amended).

# Footnotes

1)   International Vocabulary of Metrology - Basic and General Concepts and Associated Terms (VIM).

2)   OIML R134-1 "Automatic instruments for weighing road vehicles in motion and measuring axle loads –  Part 1: Metrological and technical requirements – Tests" – available to the public at www.oiml.org

3) Ministry of Transport and Communications Decree No. 341/2002 Coll., on approving technical competence, and on technical conditions for operating vehicles on the road network, as amended.

4) This government decree implements Directive 2004/22/EC of the European Parliament and of the Council of 31 March 2004 on measuring instruments into Czech legislation.