1. ------IND- 2013 0520 CZ- EN- ------ 20130925 --- --- PROJET

Managerial summary for the EC (not part of this legislation) Measuring instruments for the gas flow quantity of heating gas (either velocity or rotary piston type) which are intended for use in residential and commercial zones and light industry are placed onto the market and entered into service in the Czech Republic in accordance with Directive 2004/22/EC MID. They are subject to national metrological regulation – validation (at specified intervals). The subject of this notified legislation is simply the validation of measuring instruments which have already been put into service. This legalisation does not concern placement on the market or introduction into service (with the exception of cases pursuant to Article 23 MID). The requirements imposed upon these measuring instruments in operation are fully compatible with the MID and are based on EN 12261+A1 and EN 12480+A1 standards, which are harmonised standards for the MID. Some of the parameters specified here are properties of measuring instruments in new condition and are only noted here because preserving said properties during service is merely the subject of visual or other inspection in the scope of validation. (End of managerial summary.)

As a body with material and local competence in matters concerning the establishment of metrological and technical requirements for the specified measuring instrument and the establishment of testing methods for type approval and validation of the specified measuring instrument pursuant to Section 14 Paragraph 1 of Act No 505/1990 Coll. on metrology, as amended (hereinafter referred to as the ‘Metrology Act’), and pursuant to the provisions of Section 172 et seq. of Act No 500/2004 Coll., the Rules of Administrative Procedure (hereinafter referred to as ‘RAP’), the Czech Metrological Institute (CMI) has launched ex officio an administrative proceeding pursuant to Section 46 of the RAP on 8 July 2013, and, on the basis of the documentation, enacts the following:

I.

GENERAL MEASURE

Number: 0111-OOP-C031-13

Reference No 0313/002/13/Pos.,

establishing metrological and technical requirements for specified measuring instruments, including testing methods for validating the following specified measuring instruments:

‘instruments for measuring the gas flow quantity of heating gas (either velocity or rotary piston type) which are intended for use in residential and commercial zones and light industry’

This General Measure establishes the metrological and technical requirements for instruments for measuring the gas flow quantity of heating gas (either velocity or rotary piston type) which are intended for use in residential and commercial zones and light industry. This General Measure applies to the validation of said devices after they have been placed on the market or until they are introduced into service. The requirements of this General Measure are in accordance with the requirements of Government Regulation No 464/2005 Coll. laying down the technical requirements for measuring instruments, as amended (hereinafter referred to as the ‘Government Regulation on Measuring Instruments’)1), and make use of the relevant requirements of harmonised standards.

Instruments for measuring the gas flow quantity of heating gas (either velocity or rotary piston type) whose type was approved in accordance with the Metrology Act, as amended as of 30 October 2006, i.e. in its state before the implementation into the Czech legal order of Directive of the European Parliament and Council 2004/22/EC, as amended, are subject during validation to the same metrological requirements which were decisive in their placement on the market.

1   Basic definitions

The terms and definitions pursuant to the VIM a VIML2 as well as the terms and definitions specified in the following shall apply for the purposes of this General Measure:

1.1
Gas meter with rotary pistons (hereinafter referred to as a ‘rotary gas meter’)	
A volumetric gas meter in which a fixed measuring chamber is formed between the walls of a stationary chamber and a rotary element; each rotation cycle of the element(s) moves a defined volume of gas, which is comprehensively recorded and indicated on a display device.

1.2
Turbine gas meter	
A velocity gas flow meter in which the dynamic forces of the flowing gas cause the turbine wheel to rotate at a velocity which is the function of the volumetric flow; the number of rotations of the turbine wheel form the basis for indicating the volume which flowed through the turbine gas meter.

1.3
Ultrasound gas meter	
A velocity gas flow meter with which the measurement principle is to measure the velocity of the gas flowing in a pipe. This measurement principle is based on determining the time it takes for ultrasound signals to pass through the flowing gas between the units for transmitting and receiving the ultrasound signals.

1.4
Vortex gas flow meter	
A velocity gas flow meter which is based on an effect in which vortices are alternately separated on both sides of a non-aerodynamically shaped obstacle which gas flows around on both sides. The vortices are generated at a frequency which is proportionate to the velocity of the gas flow, and thus the size of the flow as well. A suitable sensor picks up the frequency of the vortices which are generated this way and converts them into the value of the volumetric flow in the required measurement units.

1.5
Measurement error	
A percentage value which expresses the ratio of the difference between the volume indicated by the gas meter and the volume which actually flowed through the gas meter to the actual flow volume.

2   Metrological requirements

The metrological requirements are based on the requirements of the Government Regulation on Measuring Instruments and make use of the relevant requirements of the harmonised standards.

The same metrological requirements which were decisive in the gas flow meters’ placement on the market will be applied during the validation of said meters.

2.1   Specified operating conditions

2.1.1   Flow range

The manufacturer must specify a maximum flow Qmax and a minimum flow Qmin when placing the meter on the market.

The transitional flow values Qt are specified in Table 1.

Table 1 – Transitional flow values.

| Flow range Qmin : Qmax | Transitional flow Qt |
| --- | --- |
| ( 1: 20 | 0.20Qmax |
| 1: 30 | 0.15Qmax |
| 1: 50 | 0.10Qmax |
| ( 1: 50 | 0.05Qmax |

2.1.2   Maximum operating pressure

The manufacturer shall specify the gas meter’s maximum operating pressure, which must be indicated on the gas meter.

2.1.3   Temperature range

All gas meters must be capable of meeting the requirements for the minimum ambient temperature range –10 °C to 40 °C.

The minimum temperature range of heating gas is 40 °C.

2.2   Maximum permissible error

2.2.1   Maximum permissible error during validation

The maximum permissible errors for rotary and velocity gas meters are generally specified in Table 2.

Table 2 – Maximum permissible errors

| Flow (m3/h) | Maximum permissible errors |
| --- | --- |
|  | During validation | During use |
| Qmin ( Q ( Qt | ±2.0 % | ±4 % |
| Qt ( Q ( Qmax | ±1.0 % | ±2 % |

Rotary and velocity gas flow meters in accuracy class 1.0 are given the maximum permissible error during validation in the upper flow region Qt ( Q ( Qmax.

The maximum permissible errors for household ultrasound gas meters are generally specified in Table 3.

Table 3 – Maximum permissible errors for household ultrasound gas meters

| Flow (m3/h) | Maximum permissible errors for household ultrasound gas meters |
| --- | --- |
|  | During validation | In use |
| Qmin ( Q ( Qt | ±3.0 % | ±6 % |
| Qt ( Q ( Qmax | ±1,5 % | ±3 % |

Household ultrasound gas meters in accuracy class 1.5 are given the maximum permissible error during validation in the upper flow region Qt ( Q ( Qmax.

With certain types of gas meters with which the national technical standards were applied during their approval, the maximum permissible error which was valid at the time of approval may differ from the maximum permissible error specified above; in such case, the maximum permissible error specified in the type approval certificate shall take precedence.

2.2.2   Maximum permissible error during use

In the case of measuring instruments which were validated while in use in accordance with Section 11 Paragraph 4 of the Metrology Act at the request of a party whose interests may be significantly compromised by improper measurement, the ‘maximum permissible error during use’ specified in Table 2 or Table 3 in Article 2.2.1. shall apply depending on the principle by which the gas meter was designed.

For previously used measuring instruments with which the national technical standards were applied during their approval, the maximum permissible error shall apply at two times the maximum permissible error specified in the type approval certificate.

2.3   Loss of pressure

The manufacturer must declare the maximum loss of pressure Qmax.

The mean loss of pressure for a gas meter at maximum air flow Qmax with an air density of 1.2 kg/m3 may not exceed the manufacture’s specified values or the values specified in the type approval certificate.

3   Technical requirements

The metrological requirements are based on the requirements of the Government Regulation on Measuring Instruments and make use of the relevant requirements of the harmonised standards.

The same metrological requirements which were decisive in the placement of the gas flow meter on the market will be applied during the validation of said meters.

3.1   Measuring instrument design

Rotary and velocity gas flow meters have two connection points for flanged or threaded measuring instrument connections, or one coaxial connection.

3.2   Counter

The gas meter must be equipped with a metrologically inspected counter (display).

The counter must function properly under normal conditions of use throughout the typical service life of the gas meter being tested.

If it is a cylindrical mechanical counter, it must be non-volatile and impossible to reset to zero. The rotation of the cylinder during the final tenth of its course, i.e. from 9 to 0, must cause the cylinder of the next higher digit place to move by one unit.

With electronic counters, the number of the next higher digit place must increase by one when any number changes from 9 to 0.

It must be possible to clearly and correctly read the value indicated on the counter.

3.3   Materials

The components of the gas meter in the gas flow must be resistant to the effects of the constituents of the gas medium in which they operate, and they must be resistant to wear and corrosion.

3.4   Operating position of the gas meter

The operating position in which the gas meter is to be used must be indicated on the label of the gas meter, or it must be specified on the type approval certificate.

3.5   Rectifier tube

A gas meter with which the metrological properties depend on the length of the rectifier tube connected to the gas meter must be validated and used with the rectifier tube specified by the manufacturer or on the type approval certificate.

3.6   Software

If the evaluation or indication unit contains software which is essential for the metrological properties, it must be identified and secured. The gas meter must enable the software to be identified by simple means.

3.7   Resistance to improper handling

The gas meter must be designed so that any kind of mechanical action which is capable of affecting measurement accuracy will visibly and permanently damage the measurement device or the official or security seals.

3.8   Storage and transport

Wherever possible, the connections of the gas meter must be equipped with suitable unsealed caps or covers which prevent foreign matter from getting into the meter during transport and storage.

4   Measuring instrument markings

4.1   Every gas meter must be marked with at least the following information on the counter or on a separate label:

- identification mark or name of the manufacturer,

- serial number and year of production,

- maximum flow Qmax (m3/h),

- minimum flow Qmin (m3/h),

- maximum operating pressure pmax,

- with rotary gas meters, the rated cyclic volume value V (dm3),

- the impulse number value, which corresponds with the specified recorded flow volume at the place where the impulse transmitter is to be used,

- accuracy class,

- any potential other marks, if stated in a binding manner, such as type number or trial certificate number, mark of conformity, supplementary metrological markings, etc.

Connections for pressure and temperature discharge from the gas meter must be properly marked with the appropriate labels.

The markings must be at a clearly visible location and must be permanent under normal conditions of gas meter use.

With certain gas meters whose type was approved in accordance with the national technical standards, the requirement on the measuring instrument’s markings which were valid at the time the measuring instrument received its type approval may differ from the above-mentioned official requirement. In such case, the requirements specified on the type approval certificate shall be decisive.

4.2   Gas meter with two connections

Gas meters with two connections must have the gas flow direction clearly and permanently marked by an arrow between these two connection points.

4.3   Placement of official marking

The placement of official markings is established by the type approval certificate or EC certificate of type testing or other document applied during the conformity assessment when placed on the market and into service in accordance with the Government Regulation on Measuring Instruments1).

5   Measurement device type approval

These measuring instruments are placed on the market and into service with a conformity assessment in accordance with the Government Regulation on Measuring Instruments, meaning that they are not subject to type approval.

6   Initial validation

These measuring instruments are placed on the market and into service with a conformity assessment in accordance with the Government Regulation on Measuring Instruments.

The initial validation only applies to measuring instruments which have a certificate which was valid at the time of type approval with the application of the transitional provision pursuant to Section 9 of the Government Regulation on Measuring Instruments. It also applies to measuring instruments after repairs have been conducted.

A procedure identical to the follow-up validation pursuant to Chapter 7 is applied during the initial validation.

7   Follow-up validation

7.1   Overview of tests conducted

The following tests are conducted during follow-up gas meter validations conducted after the initial validation and during validation pursuant to Section 11 Paragraph 4 of the Metrology Act:

- visual inspection,

- accuracy test,

- pressure loss test, but only in cases where this is specified on the type approval certificate.

7.2   Visual inspection

The visual inspection checks whether the gas meter submitted for validation is in conformity with the approved type or with the measuring instrument’s design, for which conformity was declared during placement on the market, and special attention must be paid to checking the correctness of the markings in the sense of Article 4.1.

The durability and legibility of the markings is also inspected.

Gas meters which have been mechanically damaged and ones which do not conform with the approved type or measuring instrument design for which conformity was declared during placement on the market shall not be tested further.

7.3   Function tests

Fulfillment of the relevant metrological and technical requirements pursuant to Chapters 2 and 3 are continuously monitored over the course of the function tests and evaluations thereof.

7.3.1   Requirements for test equipment

The test equipment must be equipped with reference calibrating devices which are suitable for testing the applicable gas meters and are valid in relation to metrology.

Pressure gauges, thermometers and reference calibrating devices for volume flow must have up-to-date metrological traceability and must state uncertainty of calibration.

The total uncertainty of establishing the actual measurement error of the gas meter must be at least three times less than the value of the maximum permissible error specified for the measuring instrument being tested.

7.3.2   Accuracy test

Gas meters are tested at the flows specified in the following Table 4.

Table 4 - Test flows for validation gas meters

| Gas meter flow ratio Qmax:Qmin | Test flows |
| --- | --- |
| 20:1 and less | Qmax; 0.7Qmax; 0.4Qmax; 0.25Qmax; 0.1Qmax; 0.05Qmax |
| Over 20:1 up to 30:1 | Qmax; 0.7Qmax; 0.4Qmax; 0.25Qmax; 0.1Qmax; 0.05Qmax; Qmin |
| Over 30:1 up to 50:1 and higher | Qmax; 0.7Qmax; 0.4Qmax; 0.25Qmax; 0.15Qmax; 0.05Qmax; Qmin |

Previous used measuring instruments which the national technical standards were applied to during their approval will be tested at the flow rates specified in the type approval certificate.

The calculated measurement error for individual test flow rates must meet the requirements for the maximum permissible error during validation specified in Table 2 or 3, or, as the case may be, in the type approval certificate.

If this is specified by the manufacturer or on type approval certificates, it will be necessary to meet the requirements for installing the gas meter, such as the length of the stabilising pipe before and after the gas meter, as well as the position of the gas meter (vertical, horizontal), during testing.

During the accuracy test, it is necessary to inspect all devices in the gas meter which are capable of generating an electronic signal commensurate to the flowing gas volume.

7.3.3   Loss of pressure test

The loss of pressure test is conducted in cases in which its value was specified by the manufacturer or is specified on the type approval certificate. Gas meters are tested using air with a density of 1.2 kg/m3 at a maximum flow of Qmax and the differential pressure is measured between input and output of the gas meter.

The mean loss of pressure, calculated as the mean measured maximum and minimum differential pressure, may not exceed the maximum permissible loss of pressure specified by the manufacturer or on the type approval certificate.

8   Declared standards

For the purposes of specifying the metrological and technical requirements for measurement devices and for the purposes of specifying testing methods during validation based on this general measure, the CMI declares Czech technical standards, other technical standards or technical documents of international or foreign organisations or other technical documents containing more detailed technical requirements (hereinafter referred to as ‘declared standards’). The CMI announces a list of these notified standards assigned to the applicable measure together with the General Measure by means accessible to the public (at the internet site www.cmi.cz).

In the scope of and under the conditions established by this General Measure, compliance with these declared standards or compliance with portions thereof is deemed to be compliance with said requirements established by the measure to which these standards (or portions thereof) apply.

II.

GROUNDS

The CMI is issuing this General Measure establishing metrological and technical requirements for the specified measuring instruments and testing methods during the validation of said specified measuring instruments for the purposes of implementing Section 24c of the Metrology Act.

Regulation No 345/2002 Coll. establishing measuring instruments for compulsory validation and measuring instruments subject to type approval, as amended, classifies measuring instruments for measuring the flow quantity of gas (either velocity or rotary piston type) as measuring instruments subject to validation under Item 1.3.10 c) in the Annex to the Second List of Specified Measuring Instruments).

The CMI hereby issues this General Measure establishing metrological and technical requirements for the specific type of measuring instruments: ‘instruments for measuring the gas flow quantity of heating gas (either velocity or rotary piston type) which are intended for use in residential and commercial zones and light industry’ and the testing methods for validating said measuring instruments for the purposes of implementing Section 24c of the Act on Metrology.

This Regulation (General Measure) was notified in accordance with European Parliament and Council Directive 98/34/EC of 22 June 1998 on the procedure for the provision of information in the field of technical standards and regulations and rules on Information Society services, as amended by Directive 98/48/EC.

III.

LEGAL INSTRUCTIONS

Legal remedies pursuant to Section 173 Paragraph 2 of the RAP may not be filed for General Measures.

In accordance with the provisions of Section 172 Paragraph 5 of the Rules of Administrative Procedure, decisions on objections may not be revoked, and appeals may not be raised against them.

The conformity of the General Measure with the statutory regulations can be evaluated in the research procedure pursuant to the provisions of Section 94 to 96 of the Rules of Administrative Procedure. The participant may submit a motion to conduct a research procedure to the administrative body which issued this General Measure. If the administrative body does not find reason to initiate a research procedure, it shall inform the submitter of this fact and substantiate its reasons within thirty days. Rulings on the initiation of a research procedure may be issued within three years as of the entry into legal effect of the General Measure in accordance with the provisions of Section 174 Paragraph 2 of the Rules of Administrative Procedure.

IV.

ENTRY INTO LEGAL EFFECT

This General Measure shall enter into legal effect on the fifteenth day as of the date of its promulgation (Section 24d of the Metrology Act).

...…………………………..............

RNDr. Pavel Klenovský

General Director

# Footnotes

1) This Government Regulation implements European Parliament and Council Directive 2004/22/EC of 31 March 2004 on Measuring Instruments into Czech legislation.

2) International Vocabulary of Metrology – Basic and general concepts and associated terms (Mezinárodní metrologický slovník – Základní a všeobecné pojmy a přidružené termíny – VIM) and the International Vocabulary of Terms in Legal Metrology (Mezinárodní slovník termínů v legální metrologii – VIML) are part of the technical harmonisation collection 'Terminology in the Field of Metrology’ (Terminologie v oblasti metrologie) which is accessible to the public at www.unmz.cz.