1. ------IND- 2013 0521 CZ- EN- ------ 20130926 --- --- PROJET

Executive summary for the EC (not part of this legislation) Fuel dispensers are placed on the market in the Czech Republic pursuant to Measuring Instruments Directive 2004/22/EC. Once they have been put into operation, they are subject to national metrology regulation – verification (by specified due dates). This notified legislation only concerns measuring instruments that have already been put into operation. It does not concern placing them on the market or putting them into operation (except for cases pursuant to Article 23 of the MID). It does not concern measuring instruments for diesel, for petrol, AdBlue or CNG. Requirements imposed upon these measuring instruments in operation are fully compatible with the MID and are based on OIML R 117, which is the harmonised standard for the MID. Some parameters listed here are properties of the measuring instrument when it is new; they are listed here because their preservation while in service is subject only to visual or other checks during verification. (End of executive summary.)

The Czech Metrology Institute (CMI), as the competent local body for stipulating metrological and technical requirements for a specified measuring instrument and for the stipulation of test methods for type approval and verification of a specified measuring instrument pursuant to § 14(1) of Act No. 505/1990 on Metrology, as amended, and pursuant to provisions of § 172 et seq. of Act No. 500/2004, the Administrative Code, as amended (hereinafter the "AC"), commenced ex officio proceedings on 22 July 2013 pursuant to § 46 of the AC, and based on supporting documents issues this:

I.

MEASURE OF A GENERAL NATURE

number: 0111-OOP-C034-13

reference no. 0313/004/13/Pos.,

stipulating metrological and technical requirements for specified measuring instruments, including test methods for verification of specified measuring instruments:

“measuring instruments and measuring systems for measuring liquefied gas flow”

This measure of a general nature stipulates metrological and technical requirements for measuring instruments and measuring systems for measuring liquefied gas flow that are applied after they are placed on the market or put into operation during their verification. These requirements are in accordance with requirements of Government Order No. 464/2005, establishing technical requirements for measuring instruments, as amended (hereinafter the “Government Order on Measuring instruments”) 1).

For measuring instruments and measuring systems for measuring liquefied gas flow whose type had been approved pursuant to the Metrology Act, i.e. prior to the implementation of Directive 2004/22/EC of the European Parliament and of the Council into Czech legislation, verification is subject to metrological requirements that were applicable to their release into circulation.

1   Basic definitions

For the purposes of this measure of a general nature, terminology and definitions contained in VIM and VIML2), Annex 7 to the Government Order on Measuring instruments and the following terminology and definitions apply:

1.1
measuring instruments and measuring systems for measuring liquefied gas flow	
measuring instruments for measuring the volume or mass flow of liquefied gases

1.2 
liquefied gas dispenser (hereinafter “LPG dispenser”)
a measuring system for fuelling motor vehicles, small boats and small airplanes with LPG (a mixture of propane and butane)

1.3
dispenser quick-release coupling
a structural part of the dispenser that represents the point where the liquid is considered dispensed or received

1.4
sales transaction
a binding relationship where the result of measurement serves as a basis for the price to be paid, where one of the parties involved in the transaction relating to measurement is the customer, and all parties accept the result of measurement obtained at a given time and place

2   Metrological requirements

Metrological requirements are based on requirements of the Government Order on Measuring Instruments and on the requirements of recommendation OIML R 117-1 - Dynamic measuring systems for liquids other than water. Part 1: Metrological and technical requirements3).

During verification, measuring instruments are subject to metrological requirements that were applicable for their release into circulation.

2.1   Specified service conditions

2.1.1   Ambient temperature range

LPG dispensers must work properly over a range of ambient temperatures specified by the manufacturer as a service temperature range, which must be at least –25 °C to 55 °C.

2.1.2   Flow range

The flow range of a measuring system specified by the manufacturer must be within the limits of the flow range of each of its parts, especially the flow meter.

The minimum required ratio Qmax : Qmin of an LPG dispenser is 5:1.

2.1.3   Measured liquid

The measured liquid must be specified by a name or type whose physical properties are specified and generally known, or must be specified directly by stating the range of relevant physical properties with which manufacturers supply these liquids, for example density or viscosity range, temperature range, or pressure range.

2.2   Measurement range

2.2.1   Flow measurement range

LPG dispensers must meet requirements for the maximum permitted error over the entire flow measurement range specified by the manufacturer.

2.2.2   Minimum measured quantity (mmq)

The minimum measured quantity of liquid the measurement of which the measuring system is intended is expressed in the form 1 × 10n, 2 × 10n or 5 × 10n volume or mass units of measurement, where n is a positive or negative whole number, or zero.

An LPG dispenser’s minimum measured quantity must not be smaller than the largest of the minimum measured quantities of any of its parts.

2.3    LPG dispenser accuracy classification

2.3.1    LPG dispenser accuracy class

For the dispensation of LPG, LPG dispensers of accuracy class 1.0 are used, engendered by the maximum permitted error of the measurement system.

2.3.2   Maximum permitted errors

2.3.2.1    Table 1 gives maximum permitted errors for volumes of two litres or more.

Table 1 – Maximum permitted errors for V ≥ 2 litres

|  | Maximum permitted error |
| --- | --- |
| Measuring systems (A) (LPG dispensers) | 1.0 % |
| Measuring instruments (B) (independent measuring instruments as parts for use in a dispenser) | 0.6 % |

## 2.3.2.2   Table 2 gives maximum permitted errors for volumes less than two litres.

Table 2 – Maximum permitted errors for V < 2 litres

| Measured volume V | Maximum permitted error |
| --- | --- |
| V < 0.1 l | 4 × the value in Table 1 used for 0.1 l |
| 0.1 l ≤ V < 0.2 l | 4 × the value in Table 1 |
| 0.2 l ≤ V < 0.4 l | 2 × the value in Table 1 used for 0.4 l |
| 0.4 l ≤ V < 1 l | 2 × the value in Table 1 |
| 1 l ≤ V < 2 l | the value in Table 1 used for 2 l |

2.3.2.3    Regardless of the measured amount, the size of the maximum permitted error is the larger of the following two values:

the absolute value of the maximum permitted error listed in Table 1 or Table 2,

the absolute value of the maximum permitted error for the smallest measured quantity Emin.

2.3.2.4    For the smallest measurements equal to two litres or more, the following conditions apply:

Condition 1:

Emin must meet the condition: Emin ≥ 2R, where R is the smallest value scale value of the indication device.

Condition 2:

Emin is given by the relationship: Emin = (2mmq) ( (A/100), where mmq is the smallest measured quantity and A is the numerical value for measuring systems (A) in Table 1.

2.3.2.5   For the smallest measured quantities smaller than two litres, condition 1 above applies and Emin is equal to double the value listed in Table 2 for measuring systems (A) from Table 1.

2.3.2.6    For a measuring system and individual measuring instruments, the maximum permitted error must not be misused, nor may one of the parties be given a systematic advantage.

2.3.3    Maximum permitted error in use

During verification of measuring instruments in use performed pursuant to § 11(4) of the Metrology Act upon request by a party whose interests could be significantly damaged by incorrect measurement, the maximum permitted error during verification pursuant to Article 2.3.2 shall apply.

2.3.4    Converted indicated datum

In the case of a converted indicated datum, the maximum permitted error for measuring systems (A) in Table 1 applies.

2.3.5    Conversion apparatus

The maximum permitted error of converted indicated data caused by the conversion apparatus is equal to ((A – B), where A and B are values listed in Table 1.

Parts of the conversion apparatus that can be tested separately:

- Counter: the maximum permitted error of indicated values of characteristic values of liquid valid for calculation, positive or negative, is equal to one tenth of the maximum permitted error for measuring systems (A) in Table 1.

Attached measuring instruments must have at least the same accuracy as given by the values in Table 3.

Table  3 – Maximum permitted errors of attached measuring instruments

| Measured value | Maximum permitted error |
| --- | --- |
| Temperature | ±0.5 ºC |
| Pressure less than MPa: from 1 to 4 MPa: greater than 4 MPa: | ±50 kPa ±5 % ±200 kPa |
| Density | ±2 kg/m3 |
| NOTE: These values apply to data for characteristic values of a liquid indicated by the conversion apparatus. |

- Calculation function accuracy: the maximum permitted calculation error for each characteristic value of the liquid, positive or negative, is equal to two fifths of the value listed under letter b) for attached measuring instruments.

3   Technical requirements

Technical requirements are based on requirements of the Government Order on Measuring Instruments and on requirements of Recommendation OIML R 117-1 Dynamic measuring systems for liquids other than water. Part 1: Metrological and technical requirements3).

During verification, measuring instruments are subject to technical requirements that were applicable for their being placed into circulation.

3.1    LPG dispenser design

A measuring system that comprises an LPG dispenser must consist of at least the following main parts:

a flow meter,

a counter,

a point of separation, usually a quick-release dispenser connector,

a differential valve,

a hydraulic circuit.

For proper operation, the following may need to be connected to in the measuring system:

a filter,

a pump,

correction apparatus.

The measuring system may be equipped with other auxiliary and supplementary devices required for proper LPG dispenser operation.

3.2    LPG dispenser design requirements

Other auxiliary devices may be part of the LPG dispenser or measuring instrument counter or they may be connected to the counter via an interface.

If only one fuelling quick-release connector can be used during dispensation, after it has been replaced on the dispenser, further dispensation must be blocked until the indicator has been zeroed.

If two or more fuelling quick-release connectors can be used simultaneously or alternately, dispensation of the liquid must be blocked after replacement of the fuelling quick-release connector until the indicator has been zeroed.

If several measuring instruments have one indicator, it must be impossible to use more than one measuring instrument at once.

It must be ensured that liquefied gas in the measuring system always remains in a liquid state. This requirement is usually met using apparatus to maintain pressure, for example a differential valve.

The temperature must be measured and an adequate temperature measurement apparatus provided.

If a device to maintain pressure is used, the system must provide means for liquid pressure measurement. The pressure gauge must be installed near the flow meter and in front of the apparatus to maintain pressure. It must be possible to verify this pressure gauge, and it must be possible to seal it after verification.

If there is an attached gas indicator, the LPG dispenser must not contain a ventilation apparatus.

A connection between the gaseous phase in the supply tank and the vehicle tank is not allowed.

In addition, in cases where the flow of liquid is stopped for emergency purposes for more than 2 minutes, dispensation of liquid must stop, and the indicator must be zeroed prior to further dispensation.

There must be a one-way valve between the gas separator and the measuring instrument. The pressure reduction caused by the one-way valve must be sufficiently small to be considered negligible.

There must be a one-way valve after the measuring instrument. The pressure reduction caused by the one-way valve must be sufficiently small to be considered negligible.

Hoses must be equipped with a special connector for “full hoses”, i.e. a self-sealing valve.

Safety features must not affect the metrological functions of the measuring system.

The design of the fuelling quick-release connector must ensure that at the moment of connection or disconnection, the amount of liquid that leaks out does not exceed the smallest specified volume deviation Emin.

If the measuring system is equipped with conversion apparatus, it must be possible to separately verify the indicated amount dispensed under conditions of measurement and a connected measuring instrument.

The dispenser may have its own supply of the measured liquid or it may be intended for installation into a central supply system.

An LPG dispenser must be equipped with a zeroing device for the volume counter and the aggregate volume counter.

If it also contains a price counter, it must be possible to zero it.

Price and volume counter zeroing devices must be arranged so that the zeroing of one automatically leads to the zeroing of the other.

It must be impossible to zero the volume and price counter during dispensation.

It must be impossible to start a new measurement if the counters are not zeroed.

If an LPG dispenser has its own supply of the measured liquid propelled by an electrical motor, it must be equipped with a device that after the electrical motor stops prevents any further dispensation of fuel until the volume counter has been zeroed.

3.3    Self-serve arrangement with LPG dispensers

Information provided by the dispenser must remain available to the parties involved in the sales transaction until the latter has been settled.

If a self-serve arrangement serves for two or more dispensers, each of them must have an identification number that is part of each piece of data that the self-serve dispenser provides. In a self-serve arrangement, data on the indication device of the dispenser, the printer and the memory device display must not differ.

The self-serve arrangement’s printer must not reproduce data from the LPG dispenser as the difference between two printed values.

The manner of payment and/or operating mode must not change prior to the end of the current measurement. The self-serve arrangement, including conditions clearly specifying operating methods, must be designed so that the customer has access to at least one primary datum from the LPG dispenser for the customer, at least until the end of the transaction, so that the customer can check the dispensed amount and price. If the self-serve arrangement summarises dispensed volumes of liquid on an on-going basis for various registered customers, the value of the scale division used for these sums must not have an effect on the smallest measured quantity.

3.4   Full service mode

If the measured data is displayed only on the LPG dispenser indicator, this indicator must be labelled with information clearly visible to the customer to the effect that this dispenser may be made available only by the operator after the current sales transaction has been completed, and that in the case of a dispute the information on the indicator is considered to be correct.

In full service mode, the transaction is settled prior to the customer leaving the dispensing location. The measurement operation ends at the moment the transaction has been settled.

3.5   Protection from unauthorised manipulation

The LPG dispenser has the required number of places for official seals.

The LPG dispenser is designed so that the measuring instrument or indicator cannot be tampered with in a way that could affect the accuracy of measurement without visible damage to an official seal.

4   Labelling

4.1    LPG dispenser labelling

Each LPG dispenser must have at least the following information on a special label:

the manufacturer’s name or brand;

the dispenser type;

the serial number and year of manufacture;

minimum flow Qmin and maximum flow Qmax;

minimum (if needed) and maximum operating pressure;

smallest measured volume;

the name or type of measured liquid, or the range of relevant characteristics;

the minimum and maximum temperature of the measured liquid;

the minimum and maximum ambient temperature;

the mechanical and electromagnetic class of the ambient environment;

the nominal AC voltage value and the limit DC voltage value;

accuracy class;

the number of the type approval certificate, if one exists, or indication of conformance;

identification of the positions of individual dispensing locations.

This label must be firmly affixed to the chassis of the dispenser in a visible location and sealed.

All information on the label on the dispenser must be legible.

4.2    Labelling of LPG dispenser parts

Each separate metrologically relevant part of an LPG dispenser, such as the flow sensor, measurement transducer, counter or indicator contains (for example on another label) the following information:

the serial number;

the manufacturer’s name or brand;

the number of the type approval certificate, if one exists;

other relevant characteristics.

All information on the label on the part must be indelible, irremovable and legible.

4.3   Indicator labelling

The indicator must contain the following information:

the unit “Litre” or the symbol “L”, “l” or “dm3” after the amount indication;

the unit of national currency (Kč) near the indication of price to be paid;

the unit, for example. “Kč/L”, “Kč/Litre” or “Kč/dm3” near the unit price indication.

Information on the minimum measured quantity must be stated near the indicator.

If the conversion function is active and the amount under basic conditions is displayed, the measurement result must be accompanied by this information, for example “volume at 15 °C”.

4.4   Recordkeeping

In the case of LPG dispensers, records of specified measuring instruments must be kept pursuant to § 18(a) of the Metrology Act in the form of record sheets. A record sheet is a document that is intended for the recording of technical information on the dispenser that is not provided on the label, primarily the serial numbers of individual parts.

Each subsystem replacement, repair, servicing or verification of the measuring instrument is recorded in this document, which is provided during dispenser verification.

5   Measuring instrument type approval

LPP dispensers are placed on the market and into operation with a conformance appraisal pursuant to the Government Order on Measuring Instruments, 2) and thus are not subject to type approval.

6.   Initial verification

These measuring instruments are placed on the market and into operation with a conformance appraisal pursuant to the Government Order on Measuring Instruments.

Initial verification applies only to measuring instruments that at the given time have a valid type approval certificate with the utilisation of transitory provisions pursuant to § 9 of the Government Order on Measuring Instruments, and to verification of measuring instruments following repairs.

The procedure used during initial verification is the same as that for subsequent verification pursuant to Chapter 7.

7   Subsequent verification

7.1   In general

During subsequent verification, i.e. each verification of the measuring instrument performed after a prior verification and during verification pursuant to § 11(4) of the Metrology Act, the following activities and tests are performed:

- a visual inspection,

- a preliminary service test,

- an accuracy test,

- a test of auxiliary and supplementary devices.

For LPG dispensers with several independent dispensing locations (multiproduct dispensers), all dispensing locations are tested separately, and these tests are evaluated individually. LPG dispensers with two ranges are first verified for the lower range and then the test is extended by a test of the highest achievable flow rate.

In the case of two flow sensors being installed in parallel, both flow sensors are first tested separately, and then a test of the entire system is performed (parallel connection) at all three stipulated flow rates.

LPG dispensers with one hydraulic measuring part and several dispensing (satellite) locations will be tested according to a regular test programme, and at the end a test at the maximum achievable flow rate will be performed from each other dispensing location (from each satellite quick-release connector).

LPG dispensers are tested with the liquid they are intended to measure.

7.2   Test equipment

Test equipment used to perform accuracy tests must have an extended uncertainty of measurement less than 1/5 of the maximum permitted error pursuant to Article 2.3.2.1.

7.3   Visual inspection

The following is assessed during a visual inspection of an LPG dispenser:

whether the measuring instrument submitted for verification and its parts conform to the approved type or with the design of a measuring instrument for which conformity has been declared during market introduction, and attention must be paid to checking labelling pursuant to Chapter 4;

whether the measuring instruments and its parts, including rubber hoses, are not physically damaged or that metal parts do not show signs of corrosion.

If the measuring instrument does not meet visual inspection requirements, no further tests are performed.

7.4   Preliminary service test

A preliminary service test involves checking that the measuring instrument is acting and working correctly, and that the counter, auxiliary devices and closure fittings are working.

7.5    LPG dispenser accuracy test

An accuracy test of an LPG dispenser and its measuring instruments is performed using the volume method, by dispensing a specified amount of liquid into a reference measuring instrument connected in series at a specified flow rate.

Test flow rates are specified according to the following relationships:

Q1 = (1 to 1.1) × Qmin,

Q2 = (0.22 to 0.28) × Qmax,

Q3 = (0.6 to 1) × Qmax.

At each of the three test flow rates, measurement errors for the tested LPG dispenser are calculated. The test is repeated at least twice for each flow rate.

The amount of liquid used during the accuracy test must be at least:

10 dm3 at a flow rate of Q1,

20 dm3 at a flow rate of Q2 ,

30 dm3 at a flow rate of Q3.

The test amount of liquid corresponds to at least the amount that flows through at the given flow rate in 1 minute.

The LPG dispenser has passed the accuracy test if none of the relative deviations of the indicated volume exceed the maximum permitted error listed in Article 2.3.2.

If all true measurement errors have the same sign, in order to meet requirements of Article 2.3.2.6, at least one of the measurement errors at a typical flow rate between 0.25Qmax and Qmax must be less than half the maximum permitted error.

7.6   Tests of auxiliary and supplementary devices

Correct functioning of auxiliary and supplementary devices is checked, if the LPG dispenser is equipped with them. The relevant test is performed only once. Tests of auxiliary and supplementary devices may be combined with the accuracy test. The following tests are performed if they are relevant to the LPG dispenser and the device is used for obligation relationships.

7.6.1   Zeroing test

The volume counter values and if applicable the price values are read after the LPG dispenser has been zeroed.

The residual indicated value after zeroing must not be greater than half the minimum specified volume deviation Emin for counters with continuous display.

Counters with discrete display must indicate zero after being zeroed, without any ambiguity.

7.6.2   Aggregate counter test

The difference between the figure on the aggregate counter and the figure on the indication device that apply to the same measurement is specified. The figure on the aggregate counter is specified at the difference between its initial and final state.

The difference between this these figures must not be more than one increment of the aggregate counter.

7.6.3   Price counter test

For LPG dispensers with a mechanical counter, the test must be performed for at least two unit prices at maximum and minimum counter rotation rates.

For LPG dispensers with an electronic counter, the test is performed for a maximum unit price. The price indicated is compared to the price calculated as the product of the indicated volume and the unit price.

During LPG dispenser verification, the test is performed only for the unit price currently in use.

The deviation of the indicated price must not be greater than the price corresponding to Emin, or the smallest unit of currency.

7.6.4   Auxiliary volume counter test

Figures indicated by auxiliary volume counters are compared with figures on the indicator, applying to the same measurement.

The figures must not differ from each other by more than one division of the scale that is the largest of the divisions of scales of devices being compared.

7.6.5   Preset test

Measurement is performed using a volume or price preset, and the difference between the preset amount and the amount indicated by the volume or price counter at the end of the measurement is determined.

For measuring instruments in prepaid or pre-ordered mode, this difference must not exceed the minimum specified volume deviation Emin or the price corresponding to this deviation.

7.5.6    Self-service facility test

A measured quantity is taken using the self-service facility and the difference between the preset amount and the amount indicated by the volume or price counter at the end of the measurement is determined.

The results must not differ from each other. For measuring instruments in prepaid or pre-ordered mode, preset requirements also apply to the self-service facility.

7.5.7   Printer test

Figures from the printer are compared with figures on the indicator that apply to the same measurement.

These figures must not differ from one another by more than one division of the indicator.

7.6.8   Attached thermometer test

If the LPG dispenser is equipped with a thermometer, during the test, the deviation of the attached thermometer at three temperatures is determined: Tmin, 15 °C and Tmax.

During a test of the thermometer as a whole, which is composed of a temperature sensor and an electronic transducer, the ascertained deviations of the attached thermometer must not exceed the value of the maximum permitted error ±0.5 °C.

During a test of the thermometer where the temperature sensor and the electronic transducer are tested separately, the ascertained deviations of the temperature sensor must not exceed the value of the maximum permitted error ±0.4 °C, and the ascertained deviations of the electronic transducer must not exceed the value of the maximum permitted error ±0.3 °C.

7.6.9   Conversion apparatus test

If the LPG dispenser is equipped with a thermometer, during the test, the accuracy of the conversion apparatus is tested at three temperatures of the test liquid: Tmin, 15 °C and Tmax, simulated using signals

The ascertained relative deviation of indicated volume Ebv at a base temperature of 15 °C must not exceed the value of the maximum permitted error ±0.5 °C.

If the measured volume Vn is simulated using measurement signals, the relative deviation Ebv in indicated volume at the base temperature must not exceed the maximum permitted error of the conversion apparatus ±0,2 %.

8   Notified standards

To specify metrological and technical requirements for measuring devices and to specify verification test methods stemming from this measure of a general nature, the CMI notifies Czech technical standards, other technical standards or technical documents of international or foreign organisations or other technical documents containing detailed technical requirements (hereinafter "notified standards"). The CMI shall publish a list of these notified standards, with their assignation to relevant measures, along with the measure of a general nature, in a manner available to the public (on the website www.cmi.cz).

Compliance with notified standards or parts thereof is considered to be, to the extent and under the conditions stipulated by a measure of a general nature, to be compliance with those requirements stipulated by this measure to which these standards or parts thereof apply.

II.

RATIONALE

The CMI issues, towards the implementation of § 24c of the Metrology Act, this measure of a general nature, stipulating metrological and technical requirements for specified measuring instruments and test methods for verification of these specified measuring instruments.

In Decree No 345/2002, specifying measuring instruments for mandatory verification and measuring instruments subject to type approval, as amended, the annex "List of Specified Measuring Device Types" classifies measuring instruments and measuring systems for measuring liquefied gas flow under item 1.3.13 as measuring instruments subject to verification.

The CMI thus issues, towards the implementation of § 24c of the Metrology Act for the measuring instrument type specified as “measuring instruments and measuring systems for measuring liquefied gas flow”, this measure of a general nature, stipulating metrological and technical requirements for measuring instruments and measuring systems for measuring liquefied gas flow for verification of these specified measuring instruments.

This regulation (Measure of a General Nature) has been notified in accordance with Directive 98/34/EC of the European Parliament and of the Council of 22 June 1998 laying down a procedure for the provision of information in the field of technical standards and regulations, as amended by Directive 98/48/EC.

III.

INSTRUCTIONS

A measure of a general nature cannot be appealed: § 173(2) of the Administrative Code.

Pursuant to § 172(5) of the Administrative Code, a decision on objections cannot be appealed.

Conformance of a measure of a general nature with legislation can be judged in review proceedings pursuant to § 94 to 96 of the Administrative Code. A participant may ask the administrative body that issued this measure of a general nature to perform a review. If the administrative body does not see any reason to commence review proceedings, it shall inform the submitter of this fact within thirty days, giving reasons. Pursuant to § 174(2) of the Administrative Code, a ruling on the commencement of review proceedings can be issued within three years of the effective date of the measure of a general nature.

IV.

EFFECTIVE DATE

This measure of a general nature takes effect on the fifteenth day following the day of its publication (§ 24d of the Metrology Act).

...…………………………..............

RNDr. Pavel Klenovský

Director General

# Footnotes

1) This government order implements Directive 2004/22/EC of the European Parliament and of the Council of 31 March 2004 on measuring instruments, as amended, into Czech legislation.

2)  The International Vocabulary of Metrology - Basic and General Concepts and Associated Terms (VIM) and the International Vocabulary of Terms in Legal Metrology (VIML) are part of the “Terminology in the Area of Metrology” technical harmonisation compendium that is publicly available at www.unmz.cz

3)   OIML R 117-1 “Dynamic measuring systems for liquids other than water, Part 1: Metrological and technical requirements” – publicly available at www.oiml.org