1. ------IND- 2013 0523 CS- EN- ------ 20130926 --- --- PROJET

Executive summary for the EC (is not part of this legislation) Graduated flasks, burettes and pipettes used to check volume are subject to national metrology regulation – verification. This notified legislation concerns only initial verification of measuring instruments prior to their being put into operation. This type of measuring instrument is not subject to type approval. Requirements applied to these measuring instruments are based on international standards: EN ISO 385, EN ISO 648, EN ISO 835 and ISO 1042. (End of executive summary.)

The Czech Metrology Institute (CMI), as the competent local body for stipulating metrological and technical requirements for a specified measuring instrument and for the stipulation of test methods for type approval and for verification of a specified measuring instrument pursuant to § 14(1) of Act No. 505/1990, on Metrology, as amended, and pursuant to provisions of § 172 et seq. of Act No. 500/2004, the Administrative Code, as amended (hereinafter the "AC"), has on 8 July 2013 commenced ex officio proceedings pursuant to § 46 of the AC, and based on supporting documents issues this:

I.

MEASURE OF A GENERAL NATURE

number: 0111-OOP-C030-13

reference no. 0313/001/13/Pos.,

stipulating metrological and technical requirements for specified measuring instruments, including test methods for verification of specified measuring instruments:

“graduated flasks, burettes and pipettes used to check volume”

1   Basic definitions

For the purposes of this measure of a general nature, terminology and definitions contained in VIM and VIML 1) and the following terminology and definitions apply:

1.1
volumetric glassware	
a glass measuring instrument calibrated for a certain volume of liquid, which is indicated with a gauge mark or scale, and is for example used to check volume for laboratory purposes

1.1.1
volumetric flask	
a glass container that serves to determine the volume of a liquid by filling it to a specified gauge mark

NOTE: Volumetric flasks are labelled with the symbol “In”.

1.1.2
pipette

a glass tube with a narrow tip for drawing in and then discharging a precise volume of liquid

NOTE: Pipettes are labelled with the symbol “Ex”.

1.1.2.1
graduated pipette	
a pipette with a scale for measuring a volume of liquid

1.1.2.2
volumetric (non-graduated) pipette	
a pipette with one mark for measuring a volume of liquid

1.1.3
burette
a glass tube with a stopcock at the bottom, marked with a scale, and intended for dispensing a precise volume of liquid

NOTE: Burettes are labelled with the symbol “Ex”.

1.2
meniscus
the curved surface of a liquid in a tube, whose shape depends on the liquid’s surface tension

1.3
nominal volume	
the largest value of volume marked on volumetric glassware

1.4
true volume	
the volume value of volumetric glassware obtained through measurement

1.5
waiting time	
the duration, specified for volumetric glass calibrated for discharged volume, starting from the end of discharge of liquid just above the graduation line until the moment the meniscus begins to be set at the graduation line, usually 15 s

1.6
discharge time	
a specified time of continuous liquid discharge from a pipette or burette, which is counted from the complete opening of the pipette’s suction opening or the opening of the burette’s stopcock to maximum flow until the end of continuous discharge

1.7
discharged volume	
the volume of liquid discharged from a pipette or burette

1.8
partial discharge	
discharge of liquid from the zero graduation mark in the upper part of the scale to any graduation mark of a graduated pipette or burette

1.9
complete discharge	
discharge of liquid from any graduation mark to the tip of a graduated pipette or burette

1.10
maximum permitted volume error	
the maximum permitted deviation from the nominal volume value of volumetric glass

2   Metrological requirements

2.1   specified service conditions

The reference ambient temperature is 20 °C and the reference pressure is atmospheric pressure.

2.2   Accuracy class

Three accuracy classes are specified:

classes A and AS for higher accuracy;

class B for lower accuracy.

For graduated volumetric glass of all accuracy classes, the maximum permitted volume error must not be greater than the volume corresponding to the smallest graduation of the scale.

2.3   Nominal volume

2.3.1    The numeric value of the nominal volume of volumetric glass is chosen from the series 10 – 20 – 25 – 50 or multiples of 10n thereof (where n is a whole negative or positive number).

2.3.2    The numeric value of the volume corresponding to the volume of the smallest graduation for graduated volumetric glass is chosen from the series 1 – 2 – 5 or multiples of 10n thereof (where n is a whole positive number).

2.3.3    The volume of measured liquid is determined by the level of the surface, or meniscus, relative to the graduation line of a scale or nominal volume mark.

For transparent liquids, the lowest point of the meniscus must be at the level of the upper edge of the mark.

For opaque liquids, the edge of the meniscus must be at the level of the upper edge of the mark.

2.4    Maximum permitted volume error

2.4.1    Maximum permitted volume error for volumetric flasks

The volume of volumetric flasks is defined as the volume of water at a temperature of 20 °C, expressed in millilitres, contained by the flask if it is filled to the upper edge of the mark while respecting Article 2.3.3.

The true volume of volumetric flasks may deviate from the nominal volume at most by the maximum permitted volume error listed in Tables 1 and 2.

Table 1 – Dimensions and maximum permitted volume errors for volumetric flasks with narrow necks

| Basic dimensions | Maximum permitted volume error |
| --- | --- |
| Nominal volume ml | Inside neck diameter mm | Minimum distance of mark from start of neck mm | Class A ml | Class B ml |
| 1 | 7 ± 1 | 5 | ±0.025 | ±0.050 |
| 2 | 7 ± 1 | 5 | ±0.025 | ±0.050 |
| 5 | 7 ± 1 | 5 | ±0.025 | ±0.050 |
| 10 | 7 ± 1 | 5 | ±0.025 | ±0.050 |
| 20 | 9 ± 1 | 5 | ±0.040 | ±0.080 |
| 25 | 9 ± 1 | 5 | ±0.040 | ±0.080 |
| 50 | 11 ± 1 | 10 | ±0.060 | ±0.120 |
| 100 | 13 ± 1 | 10 | ±0.100 | ±0.200 |
| 200 | 15.5 ± 1.5 | 10 | ±0.150 | ±0.300 |
| 250 | 15.5 ± 1.5 | 10 | ±0.150 | ±0.300 |
| 500 | 19 ± 2 | 15 | ±0.250 | ±0.500 |
| 1 000 | 23 ± 2 | 15 | ±0.400 | ±0.800 |
| 2 000 | 27.5 ± 2.5 | 15 | ±0.600 | ±1.200 |
| 5 000 | 38 ± 3 | 15 | ±1.200 | ±2.400 |

Table 2 – Dimensions and maximum permitted volume errors for volumetric flasks with wide necks

| Basic dimensions | Maximum permitted volume error |
| --- | --- |
| Nominal volume ml | Inside neck diameter mm | Minimum distance of mark from start of neck mm | Class A ml | Class B ml |
| 1 | 9 ± 1 | 5 | ±0.040 | ±0.080 |
| 10 | 9 ± 1 | 5 | ±0.040 | ±0.080 |
| 20 | 11 ± 1 | 5 | ±0.060 | ±0.120 |
| 25 | 11 ± 1 | 5 | ±0.060 | ±0.120 |
| 50 | 13 ± 1 | 10 | ±0.100 | ±0.200 |
| 1 000 | 27.5 ± 2.5 | 15 | ±0.600 | ±1.200 |

2.4.2    Maximum permitted volume error for graduated pipettes

The volume of graduated pipettes is defined as the volume of water at a temperature of 20 °C, expressed in millilitres, discharged from a graduated pipette from the zero graduation line to any graduation line of individual volumes towards the discharge tip.

Discharge volume errors must not exceed the maximum permitted error values listed in Table 3. These limits represent the maximum permitted volume errors at any point of the scale and also the maximum permitted difference between volume errors at any two points.

Table 3 – Volumes, graduations and maximum permitted volume errors for graduated pipettes

| Nominal volume | Smallest scale graduation | Maximum permitted volume errors |
| --- | --- | --- |
|  |  | Accuracy classes A and AS | Accuracy class B |
| ml | Ml | Ml | ml |
| 0.1 | 0.01 | ±0.006 | ±0.01 |
| 0.2 | 0.01 | ±0.006 | ±0.01 |
| 0.5 | 0.01 | ±0.006 | ±0.01 |
| 1 | 0.01 | ±0.007 | ±0.01 |
| 1 | 0.1 | ±0.007 | ±0.01 |
| 2 | 0.02 | ±0.010 | ±0.02 |
| 2 | 0.10 | ±0.010 | ±0.02 |
| 5 | 0.05 | ±0.030 | ±0.05 |
| 5 | 0.10 | ±0.030 | ±0.05 |
| 10 | 0.1 | ±0.05 | ±0.1 |
| 20 | 0.1 | ±0.1 | ±0.2 |
| 25 | 0.1 | ±0.1 | ±0.2 |
| 25 | 0.2 | ±0.1 | ±0.2 |

2.4.3   Maximum permitted volume error for volumetric pipettes

The volume of volumetric pipettes is defined as the volume of water at a temperature of 20 °C, expressed in millilitres, discharged from a volumetric pipette from the nominal volume zero graduation line towards the discharge tip.

Discharge volume errors must not exceed the maximum permitted error values listed in Table 4.

Table 4 – Nominal volumes and maximum permitted volume errors

| Nominal volume | Maximum permitted volume errors |
| --- | --- |
|  | Accuracy classes A and AS | Accuracy class B |
| ml | ml | Ml |
| 0.5 | ±0.005 | ±0.010 |
| 1 | ±0.008 | ±0.015 |
| 2 | ±0.010 | ±0.02 |
| 5 | ±0.015 | ±0.03 |
| 10 | ±0.02 | ±0.04 |
| 20 | ±0.03 | ±0.06 |
| 25 | ±0.03 | ±0.06 |
| 50 | ±0.05 | ±0.10 |
| 100 | ±0.08 | ±0.15 |

2.4.4   Maximum permitted error for burettes

The volume of burettes is defined as the volume of water at a temperature of 20 °C, expressed in millilitres, discharged from a burette from any graduation line for individual volumes to the zero value in front of the discharge stopcock.

Discharge volume errors must not exceed the maximum permitted error values listed in Table 5.

Table 5 – Nominal volumes and maximum permitted volume errors

| Nominal volume | Smallest scale graduation | Maximum permitted volume errors |
| --- | --- | --- |
|  |  | Accuracy classes A and AS | Accuracy class B |
| ml | ml | ml | Ml |
| 1 | 0,01 | ±0,006 | ±0,01 |
| 2 | 0,01 | ±0,01 | ±0,02 |
| 5 | 0,01 | ±0,01 | ±0,02 |
| 5 | 0,02 | ±0,01 | ±0,02 |
| 10 | 0,02 | ±0,02 | ±0,05 |
| 10 | 0,05 | ±0,03 | ±0,05 |
| 25 | 0,05 | ±0,03 | ±0,05 |
| 25 | 0,10 | ±0,05 | ±0,10 |
| 50 | 0,10 | ±0,05 | ±0,10 |
| 100 | 0,20 | ±0,10 | ±0,20 |

3   Technical requirements

3.1   Measuring instrument design

3.1.1    Measuring instrument appearance and shape

3.1.1.1   The axis of the part of the volumetric glass with a scale must be vertical. Graduated flasks must stand on a horizontal base surface without rocking.

3.1.1.2   The discharge tip of pipettes and burettes must be solid, smooth and slightly conical, without a sudden narrowing of the opening.

3.1.1.3    The shape of pipettes and burettes must enable liquid to flow out evenly.

3.1.2    Scales and marks

3.1.2.1    Graduation lines, marks, numbering and labelling must be easy to read, permanent and resistant to conditions of use.

3.1.2.2   The graduation lines of the scale must have the same thickness, and their maximum thickness must not be greater than half the linear equivalent of the maximum permitted volume error.

On graduated volumetric glassware, the maximum thickness of a graduation line must not exceed one quarter of the minimum distance between the centres of adjacent graduation lines.

On volumetric glassware that is not graduated, the gauge mark is applied around its entire circumference; a gap not greater than 10% of the circumference is permitted.

3.1.2.3   All graduation lines on the scale must be perpendicular to the longitudinal axis of the part of the volumetric glassware with the scale. For flat-bottomed volumetric flasks, the gauge mark must be parallel to the base.

3.1.2.4   Distances between the graduation lines on a scale must not exhibit any visible irregularities.

3.1.2.5   Graduation lines on a scale must be applied to the cylindrical part of volumetric glassware at least 10 mm away from where its cross-section changes. In some cases, especially for Class B volumetric glassware, graduation lines can also be located on parts that do not have a circular cross-section or on a conical section of the volumetric glass.

On volumetric glassware with one gauge mark, the number indicating the nominal volume can be located along with other information, and need not be located beside the gauge mark.

On volumetric glassware with one main graduation line and a small number of auxiliary graduation lines, the number indicating the nominal volume can be located away from the graduated scale along with other information, if the auxiliary graduation lines are properly labelled.

On graduated products with a circular cross-section, the length of the graduation lines must differ enough to be discernible.

3.1.3    Materials used

Volumetric glassware is usually manufactured from hydrolytic-grade glass with a longitudinal thermal expansion coefficient of ( = 3,3 × 10-6 K-1 (β = 3 × ( = 9,9 × 10-6 K-1). The glassware is characterised by a high degree of resistance to the effects of water, water vapour, acids and saline solutions, and relatively good resistance to alkalis. If a different material is used, the manufacturer must declare its longitudinal thermal expansion coefficient.

The glassware must be without visible defects and with as little internal tension as possible.

3.2    Measuring instrument types and overall dimensions

Volumetric glassware must be manufactured with an evenly graduated scale or with one gauge mark indicating the volume for filling or emptying. Numbering of the graduated scale and gauge mark gives the volume in ml or cm3. Volumetric glassware is classified into several types, and must meet the requirements outlined in the following tables.

3.2.1 Volumetric flasks

The basic body shape of graduated flasks must correspond to Table 6, and basic dimensions to Table 1 or 2.

Table 6 – Flask body shape

| Nominal volume Ml | Flask body shape |
| --- | --- |
| 1 and 2 | Conical |
| 5 to 50 | Conical or round body |
| 100 to 5 000 | Round body |

3.2.2   Graduated pipettes

The basic shape and dimensions of graduated pipettes must correspond to Table 7.

Table 7 – Dimensions

|  |  | Nominal volume ml |
| --- | --- | --- |
|  |  | 0.1 | 0.2 | 0,5 | 1 | 2 | 5 | 10 | 20 | 25 | 25 |
| Smallest graduation | ml | 0.01 | 0.01 | 0.1 | 0.01 | 0.02 | 0.05 | 0.1 | 0.1 | 0.1 | 0.2 |
|  |  |  |  |  | 0.10 | 0.10 | 0.10 |  |  |  |  |
| Basic dimensions in mm |
| Distance from the zero division line to the last division line for type 1 pipettes | min. max. | 80 170 | 120 200 | 130 220 | 160 220 | 160 220 | 140 220 | 140 220 | 140 220 | 220 290 | 180 220 |
| Distance from the highest to the lowest numbered division line for type 2, 3 and 4 pipettes | min. max. | 70 160 | 110 190 | 110 220 | 140 220 | 140 220 | 140 220 | 140 220 | 140 220 | 200 290 | 160 220 |
| Distance from the highest division line to the top of the pipette | min. | 100 |
| Minimal outer diameter of a suction tube |  | 5.25 to 9 | 5.5 to 9 |
| Length of the straight part of the tube under the lowest division line | min. | 5 |

Classification of graduated pipettes by type

Type 1: Partial discharge

Graduated pipettes are calibrated for discharge of liquid from the zero division line in the upper part of the graduated pipette to any graduation line, with the nominal volume given by the lowest graduation line.

Type 2: Complete discharge

Graduated pipettes are calibrated for discharge of liquid from any graduation line to the tip, with nominal volume given by the highest graduation line.

Type 3: Complete discharge

Graduated pipettes are calibrated for discharge of liquid from the zero division line in the upper part of the graduated pipette to any graduation line, with the nominal volume given by the discharge of the liquid down to the tip.

Type 4: Blow-out

Graduated complete discharge pipettes (see type 3), where the remaining drop of liquid is expelled by blowing it out.

3.2.2.1    Class A, AS and B graduated partial discharge pipettes (type 1)

Volumes and discharge times for graduated pipettes must correspond to Table 8.

Table 8 –Discharge times for type 1 graduated pipettes

| Nominal volume ml | Smallest scale division ml | Discharge time |
| --- | --- | --- |
|  |  | Class A | Class AS | Class B |
|  |  | s (min.) | s (max.) | s (min.) | s (max.) | s (min.) | s (max.) |
| 0.1 | 0.01 | 2 | 3 | – | – | 2 | 3 |
| 0.2 | 0.01 | 2 | 4 | – | – | 2 | 4 |
| 0.5 | 0.01 | – | – | 4 | 10 | 2 | 11 |
| 1 | 0.01 | 7 | 10 | 4 | 10 | 2 | 11 |
| 1 | 0.1 | 2 | 10 | 4 | 10 | 2 | 11 |
| 2 | 0.02 | 8 | 12 | 4 | 10 | 2 | 12 |
| 2 | 0.10 | 2 | 12 | 4 | 10 | 2 | 12 |
| 5 | 0.05 | 10 | 14 | 7 | 13 | 5 | 14 |
| 5 | 0.10 | 4 | 14 | 7 | 13 | 5 | 14 |
| 10 | 0.1 | 13 | 17 | 7 | 13 | 5 | 17 |
| 20 | 0.1 | – | – | 11 | 17 | 9 | 21 |
| 25 | 0.1 | 15 | 21 | 11 | 17 | 9 | 21 |
| 25 | 0.2 | 5 | 15 | – | – | 5 | 15 |

3.2.2.2    Class A, AS and B graduated full discharge pipettes (type 2)

Volumes and discharge times for graduated pipettes must correspond to Table 9.

Table 9 – Discharge times for type 2 graduated pipettes

| Nominal volume ml | Smallest scale division ml | Discharge time |
| --- | --- | --- |
|  |  | Class A | Class A | Class A |
|  |  | s (min.) | s (max.) | s (min.) | s (max.) | s (min.) | s (max.) |
| 0.1 | 0.01 | 1 | 3 | – | – | 1 | 3 |
| 0.2 | 0.01 | 1 | 4 | – | – | 1 | 4 |
| 0.5 | 0.01 | – | – | 4 | 10 | 2 | 11 |
| 1 | 0.01 | 5 | 7 | 4 | 10 | 2 | 11 |
| 1 | 0.10 | 2 | 7 | 4 | 10 | 2 | 11 |
| 2 | 0.02 | 6 | 9 | 4 | 10 | 2 | 12 |
| 2 | 0.10 | 2 | 9 | 4 | 10 | 2 | 12 |
| 5 | 0.05 | 8 | 11 | 7 | 13 | 5 | 14 |
| 5 | 0.10 | 4 | 11 | 7 | 13 | 5 | 14 |
| 10 | 0.1 | 10 | 13 | 7 | 13 | 5 | 17 |
| 20 | 0.1 | – | – | 11 | 17 | 9 | 21 |
| 25 | 0.1 | 11 | 16 | 11 | 17 | 9 | 21 |
| 25 | 0.2 | 11 | 16 | – | – | 9 | 21 |

3.2.2.3   Class A, AS and B graduated full discharge pipettes (type 3)

The waiting time for these pipettes is specified as 5 s.

Volumes and discharge times for graduated pipettes must correspond to Table 10.

Table 10 – Discharge times for type 3 graduated pipettes

| Nominal volume ml | Smallest scale division ml | Discharge time |
| --- | --- | --- |
|  |  | Class A | Class A | Class A |
|  |  | s (min.) | s (max.) | s (min.) | s (max.) | s (min.) | s (max.) |
| 0.1 | 0.01 | 1 | 3 | – | – | 1 | 3 |
| 0.2 | 0.01 | 1 | 4 | – | – | 1 | 4 |
| 0.5 | 0.01 | - | - | 4 | 10 | 2 | 11 |
| 1 | 0.01 | 5 | 7 | 4 | 10 | 2 | 11 |
| 1 | 0.10 | 5 | 7 | 4 | 10 | 2 | 11 |
| 2 | 0.02 | 6 | 9 | 4 | 10 | 2 | 12 |
| 2 | 0.10 | 6 | 9 | 4 | 10 | 2 | 12 |
| 5 | 0.05 | 8 | 11 | 7 | 13 | 5 | 14 |
| 5 | 0.10 | 8 | 11 | 7 | 13 | 5 | 14 |
| 10 | 0.1 | 10 | 13 | 7 | 13 | 5 | 17 |
| 20 | 0.1 | – | – | 11 | 17 | 9 | 21 |
| 25 | 0.1 | 11 | 16 | 11 | 17 | 9 | 21 |
| 25 | 0.2 | 11 | 16 | – | – | 9 | 21 |

3.2.2.4   Blow-out graduated pipettes (type 4)

Volumes and discharge times for graduated pipettes must correspond to Table 11.

Table 11 – Discharge times for type 4 blow-out pipettes

| Nominal volume ml | Smallest scale division Ml | Discharge time |
| --- | --- | --- |
|  |  | Class B |
|  |  | s (min.) | s (max.) |
| 0.1 | 0.01 | 1 | 3 |
| 0.2 | 0.01 | 1 | 4 |
| 1 | 0.01 | 2 | 7 |
| 1 | 0.10 | 2 | 7 |
| 2 | 0.02 | 2 | 7 |
| 2 | 0.10 | 2 | 7 |
| 5 | 0.05 | 4 | 10 |
| 5 | 0.10 | 4 | 10 |
| 10 | 0.1 | 4 | 10 |
| 25 | 0.2 | 5 | 15 |

3.2.3   Volumetric pipettes of class A, AS and B

Volumes and discharge times for volumetric pipettes must correspond to Table 12.

Table 12 –Discharge time in seconds

| Nominal volume ml |  | 0,5 | 1 | 2 | 5 | 10 | 20 | 25 | 50 | 100 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Class A | min. | 10 | 10 | 10 | 15 | 15 | 25 | 25 | 30 | 40 |
|  | max. | 20 | 20 | 25 | 30 | 40 | 50 | 50 | 60 | 60 |
| Class AS | min. | 6 | 7 | 7 | 9 | 11 | 12 | 15 | 20 | 25 |
|  | max. | 10 | 11 | 11 | 13 | 15 | 16 | 20 | 25 | 30 |
| Class B | min. | 4 | 5 | 5 | 7 | 8 | 9 | 10 | 13 | 25 |
|  | max. | 20 | 20 | 25 | 30 | 40 | 50 | 50 | 60 | 60 |
| Difference between the marked and measured discharge time | max. | 2 | 2 | 2 | 3 | 3 | 4 | 4 | 5 | 5 |

3.2.4   Burettes

3.2.4.1    Burettes without a waiting time

Volumes, dimensions and discharge times of burettes must correspond to Tables 13 and 14.

Table 13 –Dimensional requirements

| Nominal volume | Smallest division | Scale length | Total length |
| --- | --- | --- | --- |
| ml | ml | mm (min.) | mm (max.) | mm (max.) |
| 1 | 0,01 | 150 | 200 | 575 |
| 2 | 0,01 | 200 | 300 | 650 |
| 5 | 0,01 | 480 | 600 | 820 |
| 5 | 0,02 | 250 | 600 | 800 |
| 10 | 0,02 | 480 | 600 | 820 |
| 10 | 0,05 | 250 | 450 | 620 |
| 25 | 0,05 | 480 | 600 | 820 |
| 25 | 0,10 | 250 | 450 | 620 |
| 50 | 0,10 | 500 | 600 | 820 |
| 100 | 0,20 | 550 | 650 | 870 |

Table 14 – Discharge time

| Nominal volume ml | Smallest scale division ml | Discharge time |
| --- | --- | --- |
|  |  | Accuracy class A | Accuracy class B |
|  |  | s (min.) | s (max.) | s (min.) | s (max.) |
| 1 | 0.01 | 20 | 50 | 20 | 50 |
| 2 | 0.01 | 15 | 45 | 10 | 45 |
| 5 | 0.01 | 20 | 75 | 20 | 65 |
| 5 | 0.02 | 20 | 75 | 20 | 65 |
| 10 | 0.02 | 75 | 95 | 40 | 95 |
| 10 | 0.05 | 75 | 95 | 45 | 75 |
| 25 | 0.05 | 70 | 100 | 30 | 70 |
| 25 | 0.10 | 35 | 75 | 30 | 70 |
| 50 | 0.10 | 50 | 100 | 40 | 100 |
| 100 | 0.20 | 60 | 100 | 30 | 100 |

3.2.4.2   Burettes with waiting time

The waiting time for these burettes is specified as 30 s.

Volumes, dimensions and discharge times of burettes must correspond to Tables 13 and 15.

Table 15 – Discharge time

| Nominal volume ml | Smallest scale division ml | Discharge time |
| --- | --- | --- |
|  |  | s (min.) | s (max.) |
| 2 | 0.01 | 8 | 20 |
| 5 | 0.01 | 15 | 25 |
| 5 | 0.02 | 15 | 25 |
| 10 | 0.02 | 35 | 45 |
| 10 | 0.05 | 35 | 45 |
| 25 | 0.05 | 35 | 45 |
| 25 | 0.10 | 35 | 45 |
| 50 | 0.10 | 35 | 45 |

4   Measuring instrument labelling

Volumetric glass must be labelled with the following information:

- the manufacturer;

- the nominal volume with the unit of measure;

- accuracy class A, AS or B;

- the “In” symbol indicating that the product has been calibrated to contain or the symbol “Ex” indicating that the product has been calibrated to deliver;

- a reference temperature of 20 °C;

- the maximum permitted error;

- products for which a waiting time is prescribed must be labelled for example as “Ex + 5 s”;

- other information according to user requirements (must not be at odds with points a) to h)).

5   Measuring instrument type approval

These measuring instruments are not subject to type approval pursuant to the Metrology Act.

6   Initial verification

6.1    Overview of tests performed

The process for initial verification of volumetric glassware includes the following steps:

- a visual inspection;

a check of the volumetric glassware’s dimensions;

an accuracy test; for pipettes and burettes with a discharge time test.

6.2    Visual inspection

During a visual inspection, volumetric glassware submitted for initial verification is checked whether it meets the appropriate technical requirements pursuant to Chapter 3, and attention must be paid to whether it is correctly labelled, rendition of volume markings, and how volume is marked.

Volumetric glass must not be physically damaged or deformed and must be entirely clean.

Volumetric glass that fails a visual inspection is disqualified from further testing.

6.3   Functional tests

6.3.1    Test equipment accuracy requirements

For volume and dimension tests, regular laboratory equipment is used along with the following measuring instruments:

- reference scales of accuracy class II and III with weighing capacity corresponding to the measured volume;

- reference weights of accuracy class II and III corresponding to the mass of the measured volume;

- a glass thermometer with a measurement range of  (0 – 50) °C, graduated in steps of 0.1 °C;

- a stopwatch with readout accuracy of 1 s;

- a calliper gauge for measuring length with an accuracy of 0.02 mm, a flexible steel ruler up to 300 mm;

- a magnifying glass with at least 6x magnification.

The reference scales, reference weights, barometer, thermometer and stopwatch must have a valid metrological traceability with the uncertainty given.

The total uncertainty in establishing measurement instrument error must be at least three times smaller than the value of the maximum permitted error specified for the volumetric glassware being tested.

6.3.2   Dimension check

Technical parameters and dimensions of volumetric glass are checked according to tables with the given dimensions.

Volumetric glass that fails a check of basic dimensions is disqualified from further testing.

6.3.3    Accuracy test, for pipettes and burettes with a discharge time test

The accuracy of the volume of volumetric glassware is performed using the mass (gravimetric) method for contained volume (In) or delivered volume (Ex).

6.3.3.1    Determining the volume of volumetric container glassware marked with the “In” symbol

First, the mass of an empty, degreased, clean and dried volumetric flask m1 is determined. After it has been filled with distilled water (on a level surface and without bubbles) up to the gauge mark with the meniscus correctly set to the upper edge of the gauge mark, the mass of the full volumetric flask m2 is determined. The mass of the volume of distilled water mi is determined by calculating the difference in mass mi = m2 – m1.

Weighing must be performed quickly and immediately after weighing the temperature of the distilled water, air temperature and barometric pressure are measured.

6.3.3.2    Determining the volume of volumetric delivery glassware marked with the “Ex” symbol

First, the mass of an empty, degreased, clean and dried weighing bottle m1 is determined.

Pipettes and burettes are filled with distilled water up to the gauge mark with the meniscus correctly set, the required volume is discharged into the weighing bottle and the discharge time is measured, which is given in tables for individual types of volumetric glass.

If pipettes and burettes meet the discharge time requirement, they are wetted and filled with distilled water (must not contain bubbles) up to the gauge mark of the required volume with the meniscus correctly set, the required volume is discharged into the weighing bottle, and the mass of the filled weighing bottle m2 is measured.

The mass of the volume of distilled water mi is determined by calculating the difference in mass:

mi = m2 – m1.

Weighing must be performed quickly and immediately after weighing the temperature of the distilled water in the supply container and air temperature and barometric pressure are measured.

6.3.3.3   Processing accuracy test results

The measured values are used in the calculations below.

Volume Vi, which corresponds to the ascertained mass mi, is calculated using a correction factor Z, which includes the corresponding density of distilled water at the measured temperature during weighing, correction for air pressure and temperature, if the temperature during the test is not 20 °C and the manufacturer provides a thermal expansion coefficient β (β = 3 × () for the material out of which the volumetric glassware is made. The correction factor Z is calculated using the formula:

where		is the air density determined from mean values of air pressure and air temperature during weighing,

is the density of distilled water corresponding to the mean value of temperature ascertained during weighing,

β	is the thermal expansion coefficient of the material from which the volumetric container is made,

t	the temperature of the distilled water during weighing.

The volume Vi is calculated using the formula:

The ascertained volume of distilled water must not differ from the nominal value by more than the maximum permitted error pursuant to Article 2.4 for the corresponding type of volumetric glassware.

6.3.4    Placing an official mark on the measuring instrument

Volumetric glassware that during testing performed pursuant to this measure of a general nature has met its requirements is given an official mark by engraving, etching, sandblasting or printing.

7    Subsequent verification

Subsequent verification of volumetric flasks, burettes and pipettes used to check volume is not performed.

8   Notified standards

To specify metrological and technical requirements for measuring devices and to specify verification and type approval test methods stemming from this measure of a general nature, the CMI notifies Czech technical standards, other technical standards or technical documents of international or foreign organisations or other technical documents containing detailed technical requirements (hereinafter "notified standards"). The CMI will publish a list of these notified standards, with their assignation to relevant measures, along with the measure of a general nature, in a manner available to the public (on the website www.cmi.cz).

Compliance with notified standards or parts thereof is considered, to the extent and under the conditions stipulated by a measure of a general nature, to be compliance with those requirements stipulated by this measure to which these standards or parts thereof apply.

II.

JUSTIFICATION

The CMI issues, towards the implementation of § 24c of the Metrology Act, this measure of a general nature, stipulating metrological and technical requirements for specified measuring instruments and test methods for verification of these specified measuring instruments.

In Decree No 345/2002, specifying measuring instruments for mandatory verification and measuring instruments subject to type approval, as amended, the annex "List of Specified Measuring Device Types" classifies volumetric flasks, burettes and pipettes used to check volume under item 1.3.3 as measuring instruments subject to verification.

The CMI thus issues, towards the implementation of § 24c of the Metrology Act for volumetric flasks, burettes and pipettes used to check volume, this measure of a general nature, stipulating metrological and technical requirements for volumetric flasks, burettes and pipettes used to check volume for verification of these specified measuring instruments.

This regulation (Measure of a General Nature) has been notified in accordance with Directive 98/34/EC of the European Parliament and of the Council of 22 June 1998 laying down a procedure for the provision of information in the field of technical standards and regulations, as amended by Directive 98/48/EC.

III.

INSTRUCTIONS

A measure of a general nature cannot be appealed: § 173(2) of the Administrative Code.

Pursuant to § 172(5) of the Administrative Code, a decision on objections cannot be appealed.

Conformance of a measure of a general nature with legislation can be judged in review proceedings pursuant to § 94 to 96 of the Administrative Code. A participant may ask the administrative body that issued this measure of a general nature to carry out a review. If the administrative body does not see any reason to commence review proceedings, it shall inform the submitter of this fact within thirty days, giving reasons. Pursuant to § 174(2) of the Administrative Code, a ruling on the commencement of review proceedings can be issued within three years of the effective date of the measure of a general nature.

IV.

EFFECTIVE DATE

This measure of a general nature shall take effect on the fifteenth day following the day of its publication (§ 24d of the Metrology Act).

...……………………………..............

RNDr. Pavel Klenovský

Director General

# Footnotes

1)	The International Vocabulary of Metrology - Basic and General Concepts and Associated Terms (VIM) and the International Vocabulary of Terms in Legal Metrology (VIML) are part of the “Terminology in the Area of Metrology” technical harmonisation compendium that is publicly available at www.unmz.cz