1. ------IND- 2014 0026 CZ- EN- ------ 20140129 --- --- PROJET

Executive summary for EC (not part of this legal regulation) Measuring instruments for measuring the length of wound goods are placed on the market and put into use in the Czech Republic in accordance with Directive 2004/22/EC (Measuring Instruments Directive MID). They are subject to national metrological regulation – validation (within set deadlines). The subject of this notified regulation is only the verification of measuring instruments already in use. The regulation does not concern placement on the market or putting into use (except in the cases referred to in Article 23 of the MID). Requirements applicable to these measuring instruments in use are fully compatible with the MID Some parameters specified here are properties of a measuring instrument as new and are only set out here because maintaining them in use is subject to a visual or other control within the framework of verification. (End of executive summary.)

The Czech Metrology Institute (hereinafter "CMI"), as the authority with material and territorial jurisdiction in the laying down of metrological and technical requirements for the measuring instrument in question and the laying down of test methods for type approval and for verification of the measuring instrument in question in accordance with Section 14(1) of Act No 505/1990 on metrology, as amended (hereinafter the "Metrology Act" ), and in accordance with Section 172 et seq. of Act No 500/2004, the Code of Administrative Procedure, as amended (hereinafter the "CAP"), initiated ex officio on 22. 11. 2013 an administrative procedure pursuant to Section 46 of the CAP, and on the basis of documentation hereby issues this:

I.

GENERAL MEASURE

Number: 0111-OOP-C036-13

Reference No 0313/007/13/Pos.,

laying down the metrological and technical requirements for the measuring instruments in question, including test methods for verification of the measuring instruments in question:

‘measuring instruments for measuring the length of wound goods’

This general measure lays down the metrological and technical requirements for instruments for measuring the length of wound goods which apply in the course of their verification following placement on the market or putting into use. These requirements are in accordance with the requirements of Government Regulation No 464/2005 laying down technical requirements for measuring instruments1 ), as amended (hereinafter the "Government Regulation on Measuring Instruments").

For instruments for measuring the length of wound goods whose type was approved in accordance with the Metrology Act in the amended version effective up to 30 October 2006, i.e. prior to the implementation into Czech law of Directive 2004/22/EC of the European Parliament and of the Council, as amended, validation is subject to the metrological requirements which applied when said measuring instruments were placed onto the market, unless this General Measure states otherwise.

1   Basic concepts

For the purposes of this general measure, the terms and definitions pursuant to VIM and VIML2 ) apply, as well as the following terms and definitions

1.1
measuring instruments for measuring the length of wound goods	
serve to determine the length of any kind of wound material (such as string, cable, wire, fishing line, thread, textiles, bands of materials)

1.2
rolling meter	
measuring instrument which measures length by rolling a measuring wheel or measuring rollers which are in direct contact with the material being measured

1.2.1
metering measuring instrument	
rolling measuring instrument for incrementally metering out defined identical lengths of measured material

1.3
folding measuring instrument	
measuring instruments with which the textile is folded into sections (layers) of the same length, thereby determining the length of the textile, except the final incomplete section, which must be measured with a line length gauge

winding measuring instrument (yarn winder)	
determines length by winding the measured material on a rotating stand (such as a yarn winder) with a circumference of 1 m; the stand’s angle of rotation is the scale for length and is converted into evaluation units which indicate the number of revolutions; the resulting length of the wound material is the product of the number of revolutions and the circumference of the winder; the incomplete wind of the final revolution must be measured with a line length gauge.

machine measurement	
length of the material measured, determined by a measuring instrument and read on an evaluation unit

table measurement	
The actual length of a measured material, determined by measuring using a reference steel measuring tape

1.7
measurement error		
the difference between machine and table measurement

1.8
evaluation unit (indicator device, counter)	
device for indicating the length of the machine measurement in the relevant units of measurement or, in the case of metering measuring instruments, for indicating the number of lengths metered out (if the same length is measured, it is considered to be a repeating indicator device) or with a winder, or the number of complete sections with a folding measuring instrument

NOTE: The evaluation unit may be analogue or digital.

2   Metrological requirements

The metrological requirements are based on the requirements of the Government Order on Measuring Instruments.

For measuring instruments whose type was approved in accordance with the Metrology Act, validation is subject to the same metrological requirements which applied when said measuring instruments were placed onto the market, unless this General Measure states otherwise.

2.1   Specified operating conditions

The manufacturer of the measuring instrument establishes the operating conditions and testing is conducted at the ambient temperature in the scope of said conditions.

2.2   Maximum permissible error

2.2.1   Rolling and folding measuring instruments

Rolling and folding measuring instruments are divided into three accuracy classes according to metrological parameters. The maximum permissible errors during follow-up validation, corresponding to the individual groups, are specified in Table 1 as a percentage of the measured length; if the maximum permissible error value in per cent is lower than the multiple Lmin (the minimum measurable length established by the manufacturer) specified in the Table, the multiple Lmin shall apply as the maximum permissible error,

Table 1 – Maximum permissible errors for rolling and folding measuring instruments

| Accuracy class | Maximum permissible errors |
| --- | --- |
| I II III | ±0.125 %, but not lower than 0.005Lmin ±0.25 %, but not lower than 0.01Lmin ±0.5 %, but not lower than 0.02Lmin |

2.2.2   Winding measuring instruments

The actual effective circumference of a winder may not differ from its nominal value of Unom = 1 m by more than (6 mm.

2.3   Additional equipment

Devices for setting to zero or setting to a known initial length may not cause an error which would exceed the maximum permissible error pursuant to Table 1 at a minimum measurable length of Lmin.

The difference between the length displayed on an indicator and the length displayed on repeating indicator devices pursuant to Article 1.9 may not exceed the smallest graduation of the evaluation unit.

If printer and indicator devices are used, the length indicated and the length printed must be identical.

If other additional (supplementary) devices are used, said device may not affect the measurement accuracy.

3   Technical requirements

The technical requirements are based on the requirements of the Government Regulation on Measuring Instruments.

For measuring instruments whose type was approved in accordance with the Metrology Act, validation is subject to the same technical requirements which applied when said measuring instruments were placed onto the market, unless this General Measure states otherwise.

3.1   Winding measuring instruments

The device must ensure that measuring wheel (roller) rolls smoothly without slipping by the value corresponding to the previous length of the material measured passes through the measuring instrument smoothly, without gathering on one side.

For textiles, it holds true that the device must indicate the value of the length of the piece of material in question at the time of measuring (table measurement), as long as the textile is lying freely and uncrumpled on the horizontal surface without being stretched.

In order to eliminate errors caused by ductility when measuring textiles, the measuring instrument must ensure that the textile being measured:

- is sufficiently free when being guided to the measurement wheel (roller),

- is measured in a state in which it is not under tension,

- is removed from the measurement wheel (roller) in an even and uniform manner.

Instruments for measuring easily ductile materials must be equipped with a release device to regulate the tension of the material being measured at the point of measurement.

The device must have a fixed mark on the fixed part of the stand (frame) which serves as the initial (beginning of measurement) and end mark (end of measurement) or two separate marks, of which one serves as the initial mark and the other as the end mark.  The distance between the two marks along the material measured may be 100 mm or an entire multiple of the length, although it may not be greater than 1 m. It is acceptable for this distance to be added to the measurement result. A folding device only has the initial mark.

If the measurement result depends on the thickness, surface quality and/or dispenser type (in a large roll, for example) of the material measured, then the suitable limit must be determined by the manufacturer.

A metering measuring instrument must have a device which stops it as soon as the required length of measured material has been metered out.  This device must be adjustable for any desired length. If it is necessary to replace any kind of parts to set the length, then said parts must be marked with the measuring instrument’s serial number. The evaluating unit counts the pieces measured and may also indicate the entire length of the measured material.

3.2   Folding measuring instruments:

The correct length of the sections must be 1 m, unless specified otherwise.

Instruments for measuring textiles thicker than 1 mm must have a device for adjusting the travel of the folding bars which regulates the correct loading length (of the sections) for the given textile thickness. An instrument without such a device must be marked: ‘Only for measuring textiles up to 1 mm thick.’ If this fitting device is operated manually, there must be a safety device on the machine which only allows the machine to be started if the instrument is set to the correct thickness of the textile being measured.

The evaluation unit must count complete sections and must have a resetting device and a mark reading ‘Number of sections.’

The instrument must allow the material to move backward in the event that visual quality control of the measured material is necessary.

The starting, stopping and backward motion of the measured material must be smooth and continuous and free of grinding and impacts. In the case of instruments equipped to change speed, a test in operation must be conducted to determine the optimal speed for each material.

3.3   Display equipment

3.3.1   General

Indicator devices must provide a direct, secure, simple and unambiguous means of indicating the measured length.

- The connection between the indicator device and measuring device may not exhibit any play and may not slip.

- Indicator devices on measuring instruments must be connected to the wheels or rollers by such a means that said devices indicate a declining length in the event that the operator deliberately or as a result of an incorrect action moves the object being measured backward or uses the instrument in the opposite direction of the normal direction of use.

- The indicator device must have a resetting device.

- An indicator device which records the number of pieces metered out or the number of whole sections must record the applicable numbers of pieces or sections immediately before the measurement of the piece is ended or before the section is positioned.

3.3.2   Analogue indicator devices

An analogue indicator device meets the following requirements:

- the distance between the indicator and the unit indicated must be less than 2 mm,

- rollers (elements) which record values of less than 1 m must marked with a different colour (such as red) than the other ones, or the indication unit’s viewing aperture must have a coloured border at the location of the elements with tenths of a metre, and  tenths of a metre must always be separated by a decimal point,

- dividing lines and numeric markings must be rendered in a permanent and easily noticeable manner.

3.3.3   Digitals indicator devices

The labelling of the units of measurement must be marked or built in on the right side of the numeric data.

3.3.4   Graduation of the indicator

The value of the smallest graduation on the indicator device may not be less than 1 mm.

3.4   Measuring instrument security and fraud prevention

Instruments must be adequately secured to prevent fraudulent use. Components which the user is not allowed to disassemble or adjust must be secured to prevent such activity.

4   Measuring instrument markings

4.1  Markings on the measurement device

Instruments must be fitted with a label bearing the following information:

- the manufacturer’s name or trademark,

- type mark and year of production,

- accuracy class,

- the number of the type approval certificate (if relevant),

the nature and characteristics of the measured material for which the instrument in question can be used with, if necessary,

value of the graduations on the scale,

minimum measurable length,

maximum measurement speed,

indication of the beginning and end of measurement.

If the use of the measuring instrument requires special measures, then the essential instructions must be provided in a clear and visible manner in the immediate vicinity of the indicator device.

A label bearing mandatory markings must be secured (by a seal, for instance) or must be impossible to be removed without being destroyed.

4.2   Placement of the official seal

Seal placement is specified by the type approval certificate or EC certificate of type testing or other document or implement applied in the conformity assessment during placement on the market and introduction into service.

Official seals must be placed in such a manner as to prevent the measuring instrument from being adjusted without damaging the seals and to prevent access to components which are not allowed to be accessed.

5   Measurement device type approval

Measuring instruments for measuring the length of wound goods are placed on the market and put into use with a conformity assessment in accordance with the Government Regulation on Measuring Instruments and thus are not subject to type approval.

6   Initial validation

These measuring instruments are placed on the market and put into use with a conformity assessment in accordance with the Government Regulation on Measuring Instruments.

Initial validation applies only to measuring instruments that at the given time have a valid type approval certificate with application of the transitory provisions pursuant to Section 9 of the Government Order on Measuring Instruments, and to the validation of measuring instruments following repairs.

A procedure identical to the follow-up validation pursuant to Chapter 7 is applied during the initial validation.

7   Follow-up validation

7.1   General

7.1.1   Overview of tests conducted

The following activities and tests are conducted during the follow-up validation (i.e. each validation of the measuring instrument conducted after the previous validation) and during validation pursuant to Section 11 Paragraph 4 of the Metrology Act:

visual inspection,

test of correct measuring instrument operation,

test of evaluation unit operation,

accuracy test.

7.1.2   Testing equipment

The following measuring instruments are used for the tests:

- steel measuring tape at least 10 m long, with a graduation value of 1 mm along the length of the first decimetre and with a graduation value of 0.01 m along its entire length, with a maximum permissible error of ±1 mm,

steel line measuring instrument  with a nominal length of 1 020 mm, with a graduation value of 1 mm along its entire length, with a maximum permissible error of ±0.5 mm, The initial zero mark must be formed by an edge,

thickness gauge with a minimum measurement range of up to 10 mm which enables reading with an accuracy class of at least 0.1 mm. The diameter of the measurement surfaces must be at least 25 mm,

shifting measuring instrument with a measurement range of up to at least 300 mm and a graduation of 0.1 mm,

roll-up measuring instrument, such as a steel measuring tape with an increment value of 1 mm.

All measuring instruments must have valid metrological traceability.

Furthermore, the mandatory furnishings include a measuring table for manual length management with a length of at least 3.5 m and a width corresponding to the minimum width of the measured material. The top of the measuring table must be smooth, even and manufactured in a manner which prevents electrostatic charges from forming,

7.1.3   Testing conditions

Instruments with added equipment are tested as a whole unit.

The measuring error of a measuring instrument which has various different speeds is determined at two speeds, and at the maximum and minimum operating speed.

The measuring error of a measuring instrument for unlimited uses is determined by measuring three materials of varying thickness, surface and flexibility.

The measuring error of a measuring instrument for limited uses is determined when measuring the material for which the instrument is intended.

The measurement error of instruments intended for measuring double-layer textiles is to be determined when measuring double-layer textiles.

Textiles intended for testing instruments must be in a free state and must remain in the same environment in which the testing instrument is located for at least 24 hours.

The lengths of the pieces of textile must be at least 20 m for rolling measuring instruments and 30 m for folding measuring instruments.

The length of a textile (as well as materials of similar characteristics) intended for testing instruments is to be measured manually on a table. The length of a textile metered out with a measuring tape (a length of max. 4m is recommended) which is deemed to be the textile’s actual length is to be marked using a pin, stitch or other similar means.  The marking is to be applied at least 20 cm from the longitudinal edge of the textile. The results of the measurement on the table are used when determining the measuring instrument’s measurement error (table measurement).

The time interval between the completion of the measuring of the textile on the table and the beginning of measurement on the instrument may be no longer than 10 minutes.

7.2   Visual inspection

The visual inspection of a measuring instrument for measuring the length of wound goods is to ascertain the following:

- whether the instrument conforms to the approved type, the type certified by the EC or to the design for which the conformity was assessed in the scope of placement on the market,

- whether the coating of the measuring wheels has a consistent thickness over its entire circumference, is not worn or damaged and whether the measuring wheels and continuous belts are clean, free of threads, etc.,

the circumference of the measuring wheel and cylindricity of the measuring roller using a sliding or roll-up measuring instrument by measuring the circumference of the cylinder near both ends and in the middle,

- whether the material is fed in a straight line (in the case of rolling instruments in which the measuring material is fed in on a flat and even level)

- the legibility of the recorded information and its conformity with the unit’s indication (in the case of evaluation units with recording equipment),

- whether or not the evaluation unit adds the final incomplete section (in the case of folding instruments),

- whether continuous belts are tensioned so that the measured material is pressed to the measuring wheel (roller),

- the ease of operation of all moving parts,

- the correctness of the evaluation unit depending on the rotations of the measuring wheel (roller), for instance 10 rotations of the wheel or roller,

whether the resetting button on the evaluation unit is functioning correctly,

whether the measuring instrument functions correctly when running forward and backward,

the play of the transmission elements between the measuring wheel (roller) and the evaluation unit; this play may not exceed the maximum value amongst the maximum permissible errors.

7.3   Test of correct operation for instruments for measuring the length of wound goods

The proper operation of all parts of a measuring instrument is tested using measured material, and the following aspects are inspected in the process:

- for all types of instruments:

uniformity of operation and removal of the measured material at the place of measuring,

lateral displacement of the edges of the measured material,

the release device’s operation for regulating the tension at the place of measuring,

operation of the blocking and guard equipment,

operation of the doubling device,

operation of the device for regulating the speed of the material measured,

with rolling instruments:

operation of the measuring instrument before achieving working speed and at working speed, when stopping and changing the direction of the measured material’s movement,

the operation of the metering device by means of measuring a predetermined length of measured material, as well as the operation of the stopping device,

operation of the marking device,

with folding instruments:

operation of the feeding device on the folding bars,

operation of the folding bars and textile catching retainers

with winders:

frictional resistance when the strand is being fed; this must be negligible so that the measured material will not tear.

7.4   Test of evaluation unit operation

The operation of the evaluation unit is tested with measured material, and the following aspects are visually ascertained at the same time:

- for all types of instruments:

the evaluation unit’s interaction with the measuring wheel (roller),

the evaluation unit’s interaction with the recording equipment,

the operation of the evaluation unit when the direction of the measured material’s movement is changed,

the operation of the metering device using a predetermined length of measured material,

the counting of the pieces of a predetermined length which are metered out and the overall length of the measured material,

with folding instruments:

the conformity of the evaluation unit’s data with the actual number of sections measured,

with winders:

the conformity of the evaluation unit’s data with a set number of winder revolutions, which is ascertained by comparing the set revolutions (such as 100 revolutions) with the revolutions measured on the winder (manual rotation) or by using any kind of additional counter capable of reading off the number of revolutions of the winder in order to inspect the reading of the evaluation unit in the winder.

7.5   Accuracy test for instruments for measuring the length of wound goods

7.5.1   Accuracy test for rolling instruments

- the length of the piece of measured material is determined via table measurement,

- machine measurement of the measured material is conducted twice,

- table measurement is used to determine the actual length of the measured material.

The result is the average value from the measuring on the table Lp (table measurement) and the measuring with the measuring instrument Lm (machine measurement) for each piece of material. When testing instruments equipped with a release device, one piece must be made of ductile material.

The indication error of a measuring instrument for unlimited uses is determined using 3 pieces of material of at least two different thicknesses.

The number of a rolling metering instrument’s settings of a predetermined length which are to be tested is determined based on the number of possible settings pursuant to Table 2. One of these measurements must be made at the smallest length which can be set. Each setting is to be tested at three different speeds: minimum, intermediate and maximum. The measuring instrument’s measurement error is determined by measuring at least 10 pieces of measured material. In the case of an evaluation unit which, in addition to the individual lengths of the measured pieces of material, also indicates the overall length of the measured material, the measuring instrument’s error may not exceed the values specified in Table 1.

Table 2 – Number of prescribed settings of a predetermined length when testing metering measuring instruments

| Number of measuring instrument settings possible | Number of prescribed settings for the test |
| --- | --- |
| 1 | 1 |
| From 2 to 5 | 3 |
| From 6 to 10 | 4 |
| Over 10 | 5 |

The data from the counters and the difference between the machine and table measurement may not exceed the values specified in Table 1.

In the case of instruments with marking devices, the uniformity, legibility and regularity of the markings applied are to be tested. The error of marking placement is determined based on the length measuring results of a table measurement conducted on at least five one-metre-long sections, the length between an arbitrary marking and the beginning, and the length between the beginning and a marking on the measured material.

Measuring instruments which measure the length of a material in forward and reverse operation (such as inspection measuring instruments for indicating errors) are tested at maximum speed in both directions of motion. Before the end of the material is reached, the instrument is to be stopped and aligned precisely to the ‘End of measurement’ marking, either by hand or slow motion; the reading on the evaluation unit is then recorded and a table measurement is conducted.

The error of the measuring instrument’s reading when measuring with an interruption during forward and reverse operation is ascertained at least three times when changing the direction of movement of the material. In the process, the material is to travel at least 1 m before each measuring cycle (in forward and reverse operation).

With instruments which only measure a material’s length in forward operation and allow short sections of material to be moved back by approximately 0.5 to 1 m during measuring (such as inspection machines), the machine is to be stopped at least three times during the machine measurement cycle and backed up by approximately 0.5 m to 1 m.

A measuring instrument’s absolute reading error ΔLa is determined as the difference in the results of measuring the length of the measured material using a measuring instrument and measured on the table according to the formula:

ΔLa = Lm – Lp

The measuring instrument’s relative reading error L0 is determined according to the formula:

7.5.2   Accuracy test for folding instruments

When determining the error of folding instruments, the length of a textile is measured first on the measuring instrument and then on the table.

With folding measuring instruments, the status and operation of the device for setting the bar’s measurement operation are to be ascertained based on the thickness of the textile before the reading error is determined. The reading error of a measuring instrument with automatic bar operation is determined based on the thickness of the textile using two textiles of different thicknesses.

Measuring instruments without devices for setting the operation of the bar are to be tested based on the textile’s thickness using a textile which is 0.1 mm to 1 mm thick.

The measuring instrument’s reading error is ascertained as follows:

the instrument is stopped after two to three sections have been placed on and the reading on the evaluation unit is reset,

the end of the final section is marked as the beginning of measurement,

at least 20 sections are placed on and the instrument is then stopped,

the point where the final section folds is marked as the end of measurement,

the length of a section is measured using a steel length gauge which is carefully inserted between individual sections in the bottom, middle and top part of a stack of sections, so that the end of the measuring instrument touches the inner side of the fold in the section, and the length of the section M is read out. When measuring with folding instruments, it is necessary to check whether the length of the section M is within the limits specified in Table 3 in light of the textile’s thickness.

the textile’s length Lp is then measured on the table between the beginning and end of the measurement.

Table 3 – Lengths of sections with folding instruments

| Textile thickness (mm) | Permissible length of section M (mm) |
| --- | --- |
|  | Smallest | Greatest |
| 0.5 | 994 | 1004 |
| 1.0 | 993 | 1003 |
| 1.5 | 992 | 1002 |
| 2.0 | 991 | 1001 |
| 2.5 | 990 | 1000 |
| 3.0 | 989 | 999 |
| 3.5 | 988 | 998 |
| 4.0 | 987 | 997 |
| 4.5 | 986 | 996 |
| 5.0 | 985 | 995 |

The actual length of the section L is calculated according to the formula:

in which n is the number of sections between the beginning and end of the measurement.

The size of the dispersion of the section’s length W measured with the measuring instrument is determined as the difference between the greatest Li max and smallest Li min value of the length of ten individual sections, measured with a steel length gauge according to the formula:

W = Li max – Li min

The error in length of the individual sections ΔL is calculated according to the formula:

ΔL = Ln – L

in which Ln is the nominal section length set on the instrument.

The length of the textile, measured with the measuring instrument, is calculated according to the formula:

Lm = n Ln – C

in which:

n	is the number of sections recorded by the evaluation unit,

C	is the length of the final, incomplete section.

7.5.3   Accuracy test for a winding measuring instrument

When the test is conducted, particular attention must be paid to ensure that there is only negligible frictional resistance while the strand is being fed during operation of the strand winder, because the measured material may tear otherwise. Only the effective circumference of the top of the winder is used during the test.

The winder’s circumference U is calculated from geometric data (see Figure 1) according to the formula:

U = 3Dm + correction value (bm, hm)

in which:

Dm	is the arithmetic mean of 9 individual diameters D, measured at the centre of the winder and at a distance of 40 mm from both ends of the strand holders,

bm 	is the arithmetic mean of the widths b of six individual strand holders,

hm 	is the arithmetic mean of the curve heights  h at the round ends of six individual strand holders,

the correction value (bm, hm) is the numeric figure from Table 4.

Table 4 – Correction values

| Values in millimetres |
| --- |
| bm | 2.7 | 2.8 | 2.9 | 3.0 | 3.1 | 3.2 |
| hm |  |  |  |  |  |  |
| 0.2 | 1.21 | 1.28 | 1.35 | 1.42 | 1.50 | 1.57 |
| 0.3 | 0.90 | 0.96 | 1.02 | 1.09 | 1.15 | 1.22 |
| 0.4 | 0.70 | 0.75 | 0.80 | 0.85 | 0.91 | 0.96 |
| 0.5 | 0.61 | 0.64 | 0.68 | 0.72 | 0.76 | 0.80 |

Legend

a) 	winder with strand holders	D	winder diameter

b) 	strand holder	h	curve height

b	strand holder width

Figure 1 – Winder geometry

A repeat measurement at a minimum of 5 points evenly distributed over the entire width of the winder is conducted in order to determine the circumference of the winder U. The circumference of the winder’s roll can be determined by such means as using plastic or paper foil which is wrapped around the circumference of the roll.  Once a marking has been made (the initial marking is analogous to the final one), the foil is unwrapped on a level surface (such as a chair) and measured.

The circumference of the winder U is the arithmetic mean of the values measured at 5 points of the winder.

8   Notified standards

To specify metrological and technical requirements for measuring devices and to specify verification and type approval test methods stemming from this measure of a general nature, the CMI notifies Czech technical standards, other technical standards or technical documents of international or foreign organisations or other technical documents containing detailed technical requirements (hereinafter ‘notified standards’). The CMI announces a list of these notified standards assigned to the applicable measure together with the General Measure by means accessible to the public (at the internet site www.cmi.cz).

Compliance with notified standards or parts thereof is considered, to the extent and under the conditions stipulated by a measure of a general nature, to be compliance with those requirements stipulated by this measure to which these standards or parts thereof apply.

II.

G R O U N D S

The CMI is issuing this General Measure establishing metrological and technical requirements for the specified measuring instruments and testing methods during the validation of said specified measuring instruments for the purposes of implementing Section 24c of the Metrology Act.

Decree No 345/2002 establishing measuring instruments for compulsory validation and measuring instruments subject to type approval, as amended, classifies instruments for measuring the length of wound goods as measuring instruments which are subject to validation in the Second List of Specified Measurement Instruments in the Annex under Item 1.1.3.

The CMI hereby issues this General Measure establishing the metrological and technical requirements for instruments for measuring the length of wound goods and testing methods during the validation of said specified measuring instruments for the specific measuring instrument type ‘instruments for measuring the length of wound goods’ for the purpose of implementing Section 24c of Act No 505/1990 on metrology, as amended.

This legislation (General Measure) was notified in accordance with European Parliament and Council Directive 98/34/EC of 22 June 1998 on the procedure for the provision of information in the field of technical standards and regulations and rules on Information Society services, as amended.

III.

LEGAL INSTRUCTIONS

Legal remedies pursuant to Section 173 Paragraph 2 of the CAP may not be filed for General Measures.

In accordance with the provisions of Section 172 Paragraph 5 of the Code of Administrative Procedure, decisions on objections may not be revoked, and appeals may not be raised against them.

The General Measure’s conformity with the statutory regulations can be evaluated in the research procedure pursuant to the provisions of Section 94 to 96 of the Code of Administrative Procedure. The party involved may submit a motion to conduct a research procedure to the administrative body which issued this General Measure. If the administrative body does not find reason to initiate a research procedure, it shall inform the submitter of this fact and substantiate its reasons within thirty days. Rulings on the initiation of a research procedure may be issued within three years as of the General Measure’s entry into legal effect in accordance with the provisions of Section 174 Paragraph 2 of the Code of Administrative Procedure.

IV.

ENTRY INTO LEGAL EFFECT

This General Measure shall enter into legal effect on the fifteenth day following the date of its promulgation (Section 24d of the Metrology Act).

...…………………………..............

RNDr. Pavel Klenovský

General Director

# Footnotes

1 ) This Government Regulation implements Directive 2004/22 /EC of the European Parliament and of the Council of 31 March 2004 on measuring instruments, as amended, into Czech legislation.

2 ) International Vocabulary of Metrology  - Basic and General Concepts and Associated Terms (VIM) and the International Vocabulary of Terms in Legal Metrology (VIML) form part of the technical harmonisation volume "Terminology in the Field of Metrology "publicly accessible at www.unmz.cz