1. ------IND- 2014 0037 CZ- EN- ------ 20140203 --- --- PROJET

Executive summary for the EC (not part of this legislation) Instruments for measuring water flow quantity – water meters which are intended for use in residential and commercial zones and in light industry are placed on the market and introduced into service in accordance with Directive 2004/22/EC (the Measuring Instruments Directive – MID). They are subject to national metrological regulation – validation (at specified intervals). The subject of this notified legislation is simply the validation of measuring instruments which have already been put into service. This legalisation does not concern placement on the market or introduction into service (with the exception of cases pursuant to Article 23 MID). The requirements imposed upon these measuring instruments in service are fully compatible with the MID and are based on EN 12261+A1 and EN 12480+A1 standards, which are harmonised standards for the MID. Some of the parameters specified here are properties of measuring instruments in new condition and are only noted here because preserving said properties during service is merely the subject of a visual inspection or other inspection during the validation process. (End of executive summary.)

As a body with material and local competence in matters concerning the establishment of metrological and technical requirements for the specified measuring instrument and the establishment of testing methods for type approval and validation of the specified measuring instrument pursuant to Section 14 Paragraph 1 of Act No 505/1990 on metrology, as amended (hereinafter referred to as the ‘Metrology Act’), and pursuant to the provisions of Section 172 et seq. of Act No 500/2004, the Rules of Administrative Procedure (hereinafter referred to as ‘RAP’), the Czech Metrological Institute (CMI) has launched ex officio an administrative proceeding pursuant to Section 46 of the RAP on 4 July 2013, and, on the basis of the documentation, enacts the following:

I.

G E N E R A L   M E A S U R E

Number: 0111-OOP-C035-13

Reference No 0313/006/13/Pos.,

laying down metrological and technical requirements for specified measuring instruments, including the testing methods for the validation of the following specified instruments:

‘instruments for measuring water flow quantity – water meters which are intended for use in residential and commercial zones and in light industry’

This General Measure establishes the metrological and technical requirements for instruments for measuring water flow quantity, specifically water meters which are intended for use in residential and commercial zones and in light industry; said requirements apply during the process of validating said measuring instruments after they have been placed onto the market or introduced into service. These requirements are in accordance with the requirements of Government Order No 464/2005, establishing technical requirements for measuring instruments, as amended (hereinafter the ‘Government Order on Measuring Instruments’) as well as with the requirements of the harmonised standards.

For water meters with type approval under the Metrology Act, in the amended version effective up to 30 October 2006, i.e. prior to the implementation of Directive 2004/22/EC of the European Parliament and of the Council, as amended, into Czech law, validation is subject to the metrological requirements that were applicable when said water meters were placed into circulation.

1   Basic definitions

The terms and definitions pursuant to the VIM a VIML as well as the terms and definitions specified in the following shall apply for the purposes of this General Measure:

1.1
measuring instrument for water flow quantity, water meter	
instrument intended for the continuous measurement, recording and display of the volume of water flowing through it under defined operating conditions

1.2
flow
proportion of the actual volume of water which flowed through the water meter and the time during which said volume flowed through the water meter; expressed in m3/h

1.3
minimum flow rate Q1 	
the minimum flow rate at which the water meter is required to operate within the limits of the maximum permissible error

NOTE Minimum flow rate was previously noted as qmin or Qmin (which is why provisions concerning Q1 also relate accordingly to qmin or Qmin).

1.4
transitional flow rate Q2	
flow rate between the constant flow rate Q3 and the minimum flow rate Q1, and which divides the flow range into two regions: the ‘lower region’ and the ‘upper region,’ each of which is characterised by its own maximum permissible error

NOTE The transitional flow rate was previously noted as qt or Qt (which is why provisions concerning Q2 also relate accordingly to qt or Qt).

1.5
constant flow rate Q3	
the maximum flow rate at which the water meter is required to operate under the defined operation conditions by a suitable means within the limits of the maximum permissible error

NOTE The constant flow rate was previously noted as qp or Qp, or also qn (which is why provisions concerning Q3 also relate accordingly to qp, Qp or qn).

1.6
overload flow rate Q4 	
the maximum flow rate at which the water meter is required to operate for a short period of time within the limits of its maximum permissible error while retaining its metrological function, and consequently operating within the limits of the maximum permissible error

NOTE The overload flow was previously noted as qs or Qs (which is why provisions concerning Q4 also relate accordingly to qs or Qs).

1.7
test flow	
average flow during a test, calculated based on an indication from a calibrated reference device; proportion of the actual volume of water which flowed through the water meter and the time during which said volume flowed through the water meter

1.8
nominal diameter (DN)	
alphanumeric designation for the size of the part of pipe systems which is used for reference purposes; it consists of the letters DN followed by a whole number, which is indirectly related to the size of the aperture or outside diameter of the end connections in mm

1.9
statistical selection test	
test in which the decision whether or not the basic set meets the requirements based on the selection test

1.10
basic set	
set of water meters from which the selection is made and which the results of the statistical selection test are measured by

1.11
selection
set of water meters which are selected from the basic set and submitted for testing for the decision on the basic set

2   Metrological requirements

The metrological requirements for water meters intended to be used in residential and commercial zones and in light industry are based on the requirements of the Government Regulation on Measuring Instruments and use the relevant requirements from the harmonised standards.

The same metrological requirements which were decisive in the placement of the water meter on the market are to be applied when validating said meters.

2.1   Rated operating conditions

A water meter’s defined operating conditions are specified by the manufacturer.

Range of the defined operating conditions:

the water temperature range must meet at least the following conditions:

from 0.1 °C to 30 °C for water meters for cold water, or

from 30 °C to 90 °C for water meters for cold water, or

it may operate within both ranges.

the relative operating pressure range is from 0.3 bar to at least 10 bar.

a nominal value for the alternating current voltage source must be defined for the power source, or limit values must be defined for direct current voltage.

2.2   Measurement range

The values for the water flow rate range must meet the following conditions:

Q3/Q1 ( 10

Q2/Q1 = 1.6

Q4/Q3 = 1.25

2.3   Maximum permissible error

2.3.1    Maximum permitted error in lower flow rate region

The maximum permissible error for the volume delivered at a flow rate between the minimum flow rate Q1 (inclusive) and the transitional flow rate Q2 (exclusive), positive or negative, is 5 % for water of any temperature.

2.3.2    Maximum permitted error in upper flow rate region

The maximum permissible error for the volume delivered at a flow rate between the transitional flow rate Q2 (inclusive) and the transitional flow rate Q4 (exclusive), positive or negative, is:

2 % for water with a temperature of ( 30 °C,

3 % for water with a temperature of ( 30 °C,

2.3.3   Maximum permissible error during use

When validating measuring instruments during use in accordance with Section 11 Paragraph 4 of the Metrology Act at the request of a party whose interests may be significantly compromised by improper measuring, then twice the maximum permissible error specified in Articles 2.3.1 and 2.3.2. shall be applied as the maximum permissible error.

2.3.4   Assessing measuring instrument error during validation

It must be impossible to misuse the maximum permissible error or abuse it to the systematic disadvantage of one party when validating a water meter.

If all errors within the measuring instrument’s measurement range have the same sign during this validation, then at least one of the errors must be less than half of the maximum permissible error.

2.4   Measuring instrument classification

2.4.1   Measuring instrument pressure classes

The minimum allowable pressure (mAP) must be 30 kPa (0.3 bar).

Measuring instruments are classified by the maximum allowable pressure (MAP) conforming to the values of the ISO series selected by the manufacturer. The water meter must be capable of withstanding the internal pressure according to the applicable pressure class and must be tested using the applicable test.

2.4.2   Measuring instrument temperature class

The measuring instruments are classified by the water temperature class conforming to the various ranges selected by the manufacturer.

2.5   Other significant metrological characteristics

2.5.1   Flow is absent

If the flow is absent, the indication of volume may not be changed.

2.5.2   Reverse flow

In the case of measuring instruments designed to measure reverse flow, the constant flow rate and measurement range may be different in each direction.

The manufacturer must specify whether the measuring instrument is designed to measure reverse flow. If it is, the reverse flow volume must be either deducted from the indicated volume, or it must be recorded separately. The maximum permissible error pursuant to Articles 2.3.1 and 2.3.2 must be complied with for both the forward and reverse flow.

3   Technical requirements

The metrological requirements for water meters intended to be used in residential and commercial zones and in light industry are based on the requirements of the Government Regulation on Measuring Instruments and use the relevant requirements from the harmonised standards.

During validation, measuring instruments are subject to the technical requirements which were decisive for said instruments’ placement on the market.

3.1    Measuring instrument size and overall dimensions

The size of a water meter is characterised by its nominal diameter (DN). Minimum threading dimensions are defined for measuring instruments with a threaded end connection.

3.2   Display device

A water meter’s display devise must be easy to read and ensure a clear and reliable visual indication of the flow volume.

The indication device may contain elements for other testing methods, such as elements for automatic testing.

The water volume must be indicated in cubic metres. The m3 sign must be specified on the meter dial or immediately next to the numeric display.

Every display device must be equipped with a validation device for clear visual testing when validating the control element.

The graduation of validation may not exceed 0.5 % of the volume corresponding to the flow volume for 1 hour and 30 minutes at minimum flow rate Q1.

3.3   Electronic device counter

The counter may be equipped with interfaces which allow additional devices to be connected. If such an interface is used, the water meter’s hardware and software must continue to function properly and it must be impossible to influence the water meter’s metrological properties.

3.4   Auxiliary devices

A water meter may contain auxiliary devices in addition to display devices. The addition of such devices, be it temporary or permanent, may not change the measuring instrument’s metrological properties.

An auxiliary device may be used to detect a flow sensor’s activity before it is displayed on the display device.

3.5   Software

Software which is essential for the metrological properties of electronic devices must be identifiable and secured. It must be possible to identify the software by simple means.

3.6   Power supply

The means of connecting water meters with electronic devices is determined by the manufacturer.

Water meters with an electronic display device must be designed so that, should the external power supply fail (AC or DC), the volume display is not lost by the measuring instrument immediately prior to the failure and that the display remains accessible for at least one hour. Data must be saved to memory at least either once per day or after each volume equivalent to 10 minutes of flow at Q3.

3.7   Materials and design

The water meter’s materials may not be adversely affected by changes in water temperature within the operating temperature limits, i.e. the materials must be resistant to internal and external corrosion, or the materials must be protected by a suitable surface treatment.

The water meter’s display device must be protected by a transparent inspection window and a suitable type of cover may be used to provide additional protection. The water meter must include a device for eliminating the effects of condensation at places where the risk of condensation arises on the inside of the water meter’s display device.

3.8   Measuring instrument security and fraud prevention

3.8.1   Mechanical protective devices

Water meters must include protective devices which must be sealed by such means that there is no way to dismantle or alter the measuring instrument without damaging the seal or protective devices once the water meter has been sealed both before and after the water meter has been properly installed.

3.8.2   Electronic sealing device

If access to the means of modifying parameters which affect the measurement results is not protected by mechanical means, access must be protected by means of a code, password, or similar, and at least the most recent parameter modification must be saved to memory. With replaceable parts and non-replaceable detachable parts, the insertion of any parts which could affect the measurement results must be restricted.

3.9   Earthing induction water meters

The induction water meter’s electrodes and the liquid being measured must be electrically connected and grounded to ensure accurate measuring and prevent galvanic corrosion. The manufacturer’s individual installation instructions for the specific measuring instrument design must be followed, although it is generally water which is earthed by this means.

3.10   Test water conductivity

The conductivity of the test water for testing induction water meters must meet the requirements in the type approval certificate, or the EC type testing certificate.

4   Water meter markings

A water meter must be clearly and visibly labelled with the information pursuant to Article 4.1 by such means that:

- all information is listed together on a single label,

- or the information is arranged on the housing, dial of the display device, identification label or cover of the measuring instrument, provided that the cover is not removable.

4.1  Markings on the measuring instrument

A water meter must be clearly and visibly labelled with information which is important for validation:

measuring unit: cubic metre,

numeric value Q3,

the Q3/Q1 ratio  (preceded with ‘R’),

the Q3/Q1, if it differs from 1,6,

the maximum allowable pressure, if it differs from 1 MPa (10 bar),

flow direction (depicted on both sides of the body – or only on one side – with an arrow indicating the flow direction so as to be easily visible under any circumstance),

the letter V or H, if the measuring instrument is only able to operate in a vertical or horizontal position,

the temperature class, if it differs from T30,

pressure loss class, if it differs from (P 63,

class of sensitivity to irregularity in the velocity field,

the manufacturer’s name or brand,

year of manufacture (last 2 digits) and serial number (which should be as close as possible to the display device),

the type approval mark according to EU legislation,

demand factor level of the climatic and mechanical environment,

EMC class,

output signals for auxiliary devices (type/level), if any,

requirements for external power source: voltage – frequency,

In cases in which the electrical power source is a battery:

in the case of a replaceable battery, the latest date upon which the battery must be changed must be marked on the measuring instrument,

in the case of a fixed internal battery, the latest date upon which the measuring instrument must be replaced must be marked on the measuring instrument,

The water meter must also be marked with a mark indicating how it was placed on the market:

type approval mark pursuant to Act No 505/1990 on metrology, as amended, or EEC type testing certificate mark pursuant to this Government Regulation,

‘CE’ mark of conformity, in the case of water meters approved prior to this Government Regulation’s entry into legal effect,

‘CE’ mark of conformity and supplementary metrological mark in the case of water meters approved in accordance with this Government Regulation.

4.2   Placement of the official seal

The placement of official seals is defined by the type approval certificate, EC type testing certificate or other document applied during the conformity assessment when the instrument was placed on the market and introduced into service.

5   Type approval of measuring instruments

Water meters which are intended for use in residential and commercial zones and light industry are placed on the market and introduced into service in accordance with the Government Regulation on Measuring Instruments1) and are thus not subject to type approval.

6   Initial validation

Water meters which are intended for use in residential and commercial zones and light industry are placed on the market and introduced into service in accordance with the conformity assessment pursuant to the Government Regulation on Measuring Instruments.

Initial validation applies only to measuring instruments which, at the given time, have a valid type approval certificate with the application of transitional provisions pursuant to Section 9 of the Government Order on Measuring Instruments, and to the validation of measuring instruments following repairs.

A procedure identical to the follow-up validation pursuant to Chapter 7 is applied during the initial validation.

7   Subsequent validation

Follow-up validation (i.e. each validation of a measuring instrument conducted after the previous validation) and validation pursuant to Section 11 Paragraph 4 of the Metrology Act (if applicable) is to be conducted in accordance with this General Measure, regardless of whether said measuring instrument was placed on the market and into circulation in accordance with the Metrology Act or in accordance with the Government Regulation on Measuring Instruments.

During validation, measuring instruments are subject to the technical requirements which were decisive for said instruments’ placement on the market.

7.1   Overview of tests conducted

The following tests and activities are to be conducted during each validation:

- visual inspection,

- function tests:

test using static pressure,

accuracy test.

7.2   Visual inspection

This visual inspection of the water meter assesses whether:

the measuring instrument submitted for validation conforms to the approved type or to the design with which conformity was declared during placement on the market, and attention must be paid to checking labelling pursuant to Chapter 4,

the measuring instrument is mechanically damaged and metallic parts exhibit traces of corrosion,

the measuring instrument’s display of volume was not lost prior to a power cut at the location of use and whether the display remained accessible in accordance with Article 3.6 (in the case of water meters with an electronic display device),

the supply batteries are in proper condition.

If the measuring instrument does not meet the visual inspection requirements, no further tests need be performed.

7.3   Function tests

7.3.1   Testing device

The expanded measurement uncertainty of the actual volume indicated by the testing device may not exceed 1/3 of the water meter’s maximum permissible error.

The ‘collection method’ is used to determine the water meter’s measurement error. With this method, the quantity of water flowing through the water meter is collected in one or more collection tanks, and its quantity is determined by volume or weighing. The measurement error is the difference of the measured volume value indicated by the test water meter minus the indication of the calibrated testing device under reference conditions.

The defined minimum volume which must flow through the water meter during the test depends on the requirements established by testing the effects of start-up and shutdown (time error), as well as on the type and design of the display device (value of the validation graduation scale).

Measuring instruments are tested either individually or in groups. In the latter case, the individual properties of the measuring instrument must be precisely defined. Measuring instruments and testing devices must be prevented from reciprocally influencing one another. If measuring instruments are tested in series, the pressure at the output of each measuring instrument must be sufficient to prevent cavitation.

7.3.2   Test using static pressure

A test whose results are capable of demonstrating leak-tightness must be conducted using pressure equivalent to at least 1.6 MAP for a period of one minute.

7.3.3   Accuracy test

Water meters’ measurement error must be determined for at least the following three flow rates:

between Q1 and 1.1Q1,

between Q2 and 1.1Q2,

between 0.9Q3 and Q3,

Additional flow rates may be specified on the type approval certificate or conformity assessment certificate.

The temperature of the test water over the course of measuring instrument validation must be as follows, depending on temperature class:

T30, T50: 	any temperature between 0.1 °C and 30 °C

T70 and T180: 	any temperature between 0.1 °C and 30 °C and (50 ± 5) °C

T30/70 and T30/180: 	(50 ± 5) °C

Errors determined at any of the above-mentioned flow rates may not exceed the maximum permissible error specified in Articles 2.3.1 and 2.3.2 under the conditions of Article 2.3.4.

7.4   Follow-up validation of volumetric water meters for cold water using statistical selection tests

7.4   For follow-up validations of volumetric water meters for cold water installed on a water supply network, a statistical selection test of the specified basic set of water meters may be used upon request in the case of volumetric water meters for cold water.

7.4   Basic set for statistical selection test

A statistical set may consist of only volumetric water meters for cold water from a single manufacturer, of the same type, nominal diameter DN, connection method, size of the constant flow rate Q3 and ratio R of flow rate Q3/Q1 (or metrological classes).

A basic set of water meters must be defined by the same operating conditions on the water supply network. Only water meters installed in an area with the same water source may be in a single basic set. This area must be confirmed by a written declaration from the water supplier (water supply company).

Only volumetric water meters for cold water which have undergone suitable validation or a conformity assessment when introduced into service in the same calendar year may be included in the basic set.

7.4.2   Statistical method used

A statistical selection test can be conducted by means of one or two selections from the basic set of volumetric water meters for cold water in accordance with recognised statistical methods. The selection conducted may contain a specified set of reserve volumetric water meters for supplementing the test selection over the course of the tests.

Logistical and other details on the performance of the statistical selection test, including inspection plans, are established by the metrological institute conducting the validation according to its internal rules.

7.4.3   Visual inspection

All volumetric water meters submitted for selection must undergo the visual inspection pursuant to Article 7.2. If a volumetric water meter from the selection does not pass the visual inspection, it can be replaced by a volumetric water meter from the set of reserve volumetric water meters until the set of reserve volumetric water meters is exhausted.

7.4.4   Accuracy test

All volumetric water meters submitted for selection must undergo an accuracy test. The accuracy test must be conducted at three test flow rates:

minimum flow rate Q1,

transitional flow rate Q2,

constant flow rate Q3.

During the accuracy test, three measurements must be conducted at each test flow rate. If a severe error, i.e. an error exceeding 15 % (for Q2 and Q3) or 24 % (for Q1), appears during any of the measurements, then a fourth measurement which will replace that measurement must be conducted immediately, after which the measurement cannot be repeated again for the flow rate in question (the maximum test programme in this case is four tests at Q1, four tests at Q2, and four tests at Q3).

The test is declared to be suitable if at least two of three results and the arithmetic average of the results of three tests at the given flow rate are within the limits of the maximum permissible errors specified in Articles 2.3.1 and 2.3.2.

7.4.4   Evaluation of the results of the statistical selection test

7.4.4.1   A controlled selection of volumetric water meters is evaluated as being suitable if the requirements for acceptance were met according to a previously accepted inspection plan for the selection check. Otherwise, the result will be deemed unsatisfactory.

7.4.4.2   If the selection check ends in an unsatisfactory result, all volumetric water meters in the basic set will be deemed unsatisfactory. The entire unsuitable basic set of water meters evaluated, including any part thereof, may no longer be validated via statistical selection test. All water meters of the basic set which have not been dismantled may be used for the term of validity of their original validity, after which each must undergo a follow-up validation.

7.5   Testing a water meter during the term of validity of its validation at the installation location

This test is conducted to check whether the water meter, in the function of a specified measuring instrument, meets the metrological and technical requirements of this General Measure, is installed properly and meets the requirements for measurement accuracy during service.

7.5.1   Tests conducted

The test of a water meter at it installation location during the term of validity of its validation consists of:

a visual inspection of the tested water meter and its installation,

a preliminary operation test,

leak-tightness test,

accuracy test.

7.5.2   Visual inspection

The visual inspection of the water meter assesses whether:

the water meter being tested conforms with the approved type or the type for which type conformity was declared,

it is properly installed and whether it is mechanically damaged,

the information on the water meter is in accordance with the type approval certificate or document on conformity assessment,

is properly secured with official seals.

In the case of water meters which did not pass the visual inspection, this fact must be recorded in the test report, including a sufficient specification of the exact form of non-conformity. The water meter test is then continued.

7.5.3   Testing equipment

The following test equipment is used at the installation location when testing the water meter:

a technically suitable reference water meter with established metrological traceability with metrological parameters corresponding to those of the test water meter, and which is fitted with a regulation valve, or

a reference measurement vessel with a regulation valve for adjusting the flow rate during the test; the vessel must have a validation graduation scale value of 0.01 L for volumes of up to and including 10 L, and 0.03 L for volumes over 10 L; the spacing between the individual graduations on the scale must be 4 mm,

a connection tube with a suitable connection for discharge from the pipe system and a vessel regulation valve,

a thermometer with established metrological traceability with an accuracy of ±1 °C.

The measurement uncertainty of the test device may not exceed 1/3 of the water meter’s maximum permissible error.

7.5.4   Leak-tightness inspection

An etalon for the flow quantity of water is connected after the test water meter to one outlet from the water supply pipe by means of a connection tube. All other outlets must be closed.

The test water meter is not manipulated, so as to retain its present status and any potential installation effect during the test.

The inspection checks the tightness of the shut-off valve and all other valves and outlets from the water supply system which, if loose, may cause the measured water to leak between the test water meter and the etalon.  It must be ensured that all water flowing through the test water meter also flows through the etalon.

The leak-tightness of the tube connection is to be controlled visually.

7.5.5   Preliminary operation test

The test is conducted for a period of at least 5 minutes at the maximum attainable flow rate from the water supply system. During the test, it is ascertained whether there are any air bubbles in the water supply system.

7.5.6   Accuracy test

The accuracy test must be conducted at a minimum flow rate of Q1 and a transitional flow rate of Q2, calculated based on the water meter’s markings, and at the maximum attainable flow rate from the water supply system, which must at least be the range of (2 ( Q2 to Q3). The test flow rates for testing the accuracy of the water meter at the installation location are as follows:

- minimum flow rate Q1 of the water meter being tested, the average flow rate value must remain within a range of  (Q1 to 1.1 ( Q1) during the test,

- transitional flow rate Q2 of the water meter being tested, the average flow rate value must remain within a range of (Q2 to 1.1 ( Q2) during the test,

- the maximum attainable flow rate, the average flow rate value must remain within a range of (2 ( Q2 to Q3) during the test.

If it is not possible to attain a flow rate of 2 ( Q2 because the water meter is unsuitable for the specific installation and the water pressure at the installation location is inadequate, then this fact must be noted in the test report.

If it is not possible to achieve a flow rate of Q1, then this fact must also be noted in the test report and the test will not be continued.

The water temperature during the course of testing a water meter for cold water, i.e. water with a temperature between (0.1 to 30) °C, and during the course of testing a water meter for hot water, i.e. water with a temperature between (30 to 90) °C, is to be measured using a thermometer with an accuracy of ±1 °C.

The defined minimum volume of water which must flow through the water meter during the test depends on the requirements established by testing the effects of start-up and shutdown (time error), as well as on the type and design of the water meter’s display device.

During the accuracy test, at least one measurement must be conducted at each test flow rate. The test is declared to be suitable if the measurement result is within the limits of the maximum permissible errors specified in Article 2.3.3.

The accuracy test may by conducted using the collection or volume method.

7.5.6.1   Collection method

With an accuracy test conducted using the collection method, the quantity of water flowing through the water meter is collected in a measurement vessel and its quantity is determined on the measurement vessel’s scale.

The measurement error is determined by comparing the volume indicated by the water meter being tested with the quantity of water which flowed indicated on the measurement vessel.

The flow rate Q1, Q2 and maximum attainable flow rate from the water supply system pursuant to Article 7.5.6 Points a) through c) is adjusted by constricting the regulation valve at the input to the measurement vessel.

7.5.6.2   Volume method

With an accuracy test conducted using the volume method, the volume which flowed through the water meter being tested is compared with the volume which flowed through the etalon for the quantity of water which flowed using an induction flow gauge (for instance). The induction flow gauge must be calibrated, including its stilling sections (at a length of at least 10D before the measuring instrument and at least 5D after the measuring instrument), with a maximum permissible error of 0.6 %. Calibration uncertainty may not exceed ±0.2 %.

The water pressure at the output from the etalon must be adjusted using an outlet tube. The pressure after the etalon must be at least 5 kPa (0.05 bar), even at a flow rate of zero (corresponding to 0.5 m of the water column).

The flow rate Q1, Q2 and maximum attainable flow rate from the pipe system pursuant to Article 7.5.6 Points a) through c) is adjusted by constricting one of the regulation valves at the output after the etalon with the stilling sections.

7.5.6.3   Evaluation of the results of the accuracy test at the installation location

If the water meter fails to pass the external inspection or accuracy test, it is deemed not to meet the defined requirements.

8   Notified standards

For the purposes of specifying the metrological and technical requirements for measuring instruments and for the purposes of specifying the testing methods for validation based on this General Measure, the CMI shall provide notification of the Czech technical standards, other technical standards or technical documents of international or foreign organisations, or other technical documents containing more detailed technical requirements (hereinafter referred to as ‘notified standards’). The CMI announces a list of these notified standards assigned to the applicable measure together with the General Measure by means accessible to the public (at the internet site www.cmi.cz).

Notified standards or parts thereof shall be deemed as complied with, within the scope of and under the conditions laid down by this General Measure, if the requirements laid down by this measure, to which these standards or parts thereof apply, are complied with.

II.

G R O U N D S

The CMI issues this General Measure laying down metrological and technical requirements for specified measuring instruments and the testing methods for the validation of these specified measuring instruments in order to implement Section 24c of the Metrology Act.

Regulation No 345/2002 establishing measuring instruments for compulsory validation and measuring instruments subject to type approval, as amended, classifies instruments for measuring the flow quantity of water – measuring instruments for cold and hot water as measuring instruments subject to validation under Item 1.3.9 a), b) in the Annex to the Second List of Specified Measuring instruments.

The CMI hereby issues this General Measure establishing metrological and technical requirements for instruments for measuring the flow quantity of water – water meters intended for use in residential and commercial zones and light industry’ and the testing methods for validating these specified measuring instrument for the purpose of implementing Section 24c of the Act on Metrology for this specific type of measuring instrument: ‘instruments for measuring the flow quantity of water – water meters intended for use in residential and commercial zones and light industry.’

This legislation (General Measure) was notified in accordance with European Parliament and Council Directive 98/34/EC of 22 June 1998 on the procedure for the provision of information in the field of technical standards and regulations and rules on Information Society services, as amended.

III.

LEGAL INSTRUCTIONS

Legal remedies pursuant to Section 173(2) of the Rules of Administrative Procedure may not be filed for General Measures.

In accordance with the provisions of Section 172(5) of the Rules of Administrative Procedure, decisions on objections may not be revoked and appeals may not be made against them.

The General Measure’s conformity with the statutory regulations can be evaluated in the research procedure pursuant to the provisions of Section 94 to 96 of the Rules of Administrative Procedure. The party involved may submit a motion to conduct a research procedure to the administrative body which issued this General Measure. If the administrative authority finds no reasons to commence the review proceedings, it shall communicate and provide grounds for this fact within thirty days. Rulings on the initiation of a research procedure may be issued within three years as of the General Measure’s entry into legal effect in accordance with the provisions of Section 174 Paragraph 2 of the Rules of Administrative Procedure.

IV.

ENTRY INTO LEGAL EFFECT

This General Measure shall enter into legal effect on the fifteenth day following the date of its promulgation (Section 24d of the Metrology Act).

...…………………………..............

RNDr. Pavel Klenovský

Director General

# Footnotes

This Government Order implements European Parliament and Council Directive 2004/22/EC of 31 March 2004 on Measuring Instruments, as amended, into Czech legislation.

International Vocabulary of Metrology – Basic and general concepts and associated terms (Mezinárodní metrologický slovník – Základní a všeobecné pojmy a přidružené termíny – VIM) and the International Vocabulary of Terms in Legal Metrology (Mezinárodní slovník termínů v legální metrologii – VIML) are part of the technical harmonisation collection ‘Terminology in the Field of Metrology’ (Terminologie v oblasti metrologie) which is accessible to the public at www.unmz.cz.